## Sheet1
In this document I report the count rate over the complete t.o.f. range for the detectors 1-198 (monitors, backscattering and frontscattering) as a function of the consecutive runs (each of 90 uAh) during the measurement of a Polyethylene sample of 0.25 mm thickness measured on the weekend 21/22 March 2015. The last columns on the right report the average and the standard deviation of the count rates for each detector. Their ratio, to be considered as a relative error (fluctuation) on each detector stability over a long period of time, is reported as a plot here on the right. Almost all the detectors are stable within a 0.5%!
Detector/run | 17713 | 17714 | 17715 | 17716 | 17717 | 17718 | 17719 | 17720 | 17721 | 17722 | 17723 | 17724 | 17725 | 17726 | 17727 | 17728 | 17729 | 17730 | 17731 | 17732 | 17733 | 17734 | 17735 | 17736 | 17737 | 17738 | 17739 | 17740 | 17741 | 17742 | 17743 | 17744 | 17745 | 17746 | 17747 | 17748 | 17749 | 17750 | 17751 | 17752 | 17753 | 17754 | 17755 | 17756 | 17757 | 17758 | 17759 | 17760 | 17761 | 17762 | 17763 | 17764 | 17765 | 17766 | 17767 | 17768 | 17769 | 17770 | 17771 | 17772 | 17773 | 17774 | 17775 | 17776 | 17777 | 17778 | 17779 | Detector | Average | S. deviation | %
1 | 17640706 | 17631120 | 17624957 | 17588558 | 17403739 | 17638456 | 17661160 | 17640372 | 17483148 | 17608730 | 17633780 | 17639242 | 17615161 | 17617858 | 17602796 | 17606217 | 17486187 | 17583816 | 17615232 | 17607310 | 17607310 | 17616217 | 17600588 | 17546993 | 17502933 | 17615514 | 17616515 | 17630359 | 17650017 | 17670436 | 17692947 | 17711627 | 17702947 | 17657389 | 17563231 | 17602388 | 17618222 | 17598352 | 17557663 | 17571260 | 17574014 | 17554333 | 17535883 | 17525937 | 17531823 | 17533135 | 17522054 | 17508136 | 17504971 | 17514722 | 17519649 | 17440646 | 17530430 | 17546115 | 17524412 | 17565468 | 17579412 | 17589882 | 17579888 | 17581727 | 17574460 | 17567657 | 17567903 | 17576332 | 17553315 | 17558775 | 17564973 | 1 | 17581903.0597 | 58144.5843 | 0.0033071
2 | 4681052 | 4680622 | 4683352 | 4671057 | 4619273 | 4678560 | 4684500 | 4681680 | 4645330 | 4678314 | 4683732 | 4685075 | 4675931 | 4679949 | 4679868 | 4673783 | 4645196 | 4672669 | 4678993 | 4679842 | 4679842 | 4675534 | 4674108 | 4657350 | 4643283 | 4679174 | 4679723 | 4678184 | 4678436 | 4681254 | 4681368 | 4683924 | 4677238 | 4680604 | 4660968 | 4680739 | 4685463 | 4683355 | 4676815 | 4683463 | 4680658 | 4675991 | 4674460 | 4672055 | 4667675 | 4670584 | 4668849 | 4667652 | 4666494 | 4665640 | 4659400 | 4636843 | 4660599 | 4667835 | 4657212 | 4669686 | 4670406 | 4670057 | 4669009 | 4671656 | 4671664 | 4670301 | 4671488 | 4670098 | 4667917 | 4665235 | 4670298 | 2 | 4671781.5672 | 12123.8089 | 0.0025951
3 | 3865571 | 3865232 | 3861591 | 3857677 | 3795657 | 3860080 | 3864771 | 3865151 | 3825339 | 3857702 | 3865021 | 3865652 | 3860201 | 3860747 | 3856382 | 3858477 | 3830182 | 3849440 | 3861897 | 3863984 | 3863984 | 3860797 | 3859784 | 3845299 | 3833825 | 3861658 | 3862691 | 3864492 | 3861588 | 3865564 | 3863846 | 3865463 | 3864514 | 3863028 | 3844818 | 3860501 | 3863231 | 3859958 | 3853830 | 3861308 | 3859473 | 3860878 | 3860389 | 3856935 | 3857582 | 3866480 | 3860051 | 3856302 | 3858986 | 3858603 | 3853753 | 3833683 | 3850301 | 3856334 | 3848546 | 3856897 | 3859513 | 3857864 | 3856918 | 3854452 | 3858283 | 3856466 | 3860978 | 3855757 | 3854055 | 3859498 | 3859664 | 3 | 3856710.0597 | 11155.0888 | 0.0028924
4 | 3627293 | 3623718 | 3626494 | 3616811 | 3563014 | 3619743 | 3627862 | 3626048 | 3589541 | 3620291 | 3629037 | 3622073 | 3619350 | 3625067 | 3623402 | 3626136 | 3594613 | 3614045 | 3624874 | 3625898 | 3625898 | 3623964 | 3621431 | 3607421 | 3595820 | 3623823 | 3627541 | 3627047 | 3623157 | 3626710 | 3626223 | 3628014 | 3628206 | 3624286 | 3605949 | 3619130 | 3624211 | 3623399 | 3615615 | 3623484 | 3622395 | 3624749 | 3618819 | 3620636 | 3618337 | 3622366 | 3621175 | 3616776 | 3620900 | 3619972 | 3615313 | 3595629 | 3614277 | 3624903 | 3614755 | 3620603 | 3623341 | 3622631 | 3619244 | 3620221 | 3623520 | 3618025 | 3618279 | 3621374 | 3620171 | 3622131 | 3618830 | 4 | 3619194.194 | 10556.0458 | 0.0029167
5 | 3571291 | 3571232 | 3574124 | 3567597 | 3511669 | 3571496 | 3577592 | 3571396 | 3538504 | 3566799 | 3570599 | 3571127 | 3569109 | 3572517 | 3569988 | 3567523 | 3539676 | 3565959 | 3573025 | 3572464 | 3572464 | 3571697 | 3571134 | 3558072 | 3544988 | 3573172 | 3573800 | 3572286 | 3570318 | 3570671 | 3573406 | 3575112 | 3574535 | 3570423 | 3555490 | 3570624 | 3578348 | 3570827 | 3562933 | 3570614 | 3571873 | 3570017 | 3571341 | 3570646 | 3565625 | 3566864 | 3570272 | 3567784 | 3565800 | 3564988 | 3561784 | 3544030 | 3562471 | 3566356 | 3554376 | 3564518 | 3569459 | 3567593 | 3565663 | 3566934 | 3565639 | 3563978 | 3569614 | 3565965 | 3567583 | 3570035 | 3570280 | 5 | 3566508.791 | 10309.2756 | 0.0028906
6 | 2694770 | 2691233 | 2692072 | 2687199 | 2650640 | 2688472 | 2696258 | 2691423 | 2667819 | 2689886 | 2693705 | 2692422 | 2690652 | 2692048 | 2691544 | 2688012 | 2666177 | 2684289 | 2688334 | 2691239 | 2691239 | 2690226 | 2690858 | 2678951 | 2670444 | 2691075 | 2691167 | 2692306 | 2687844 | 2686798 | 2686556 | 2691384 | 2689747 | 2687818 | 2673324 | 2683704 | 2690631 | 2685850 | 2684666 | 2686286 | 2686198 | 2684470 | 2686706 | 2681900 | 2681517 | 2683782 | 2681230 | 2679555 | 2682367 | 2679625 | 2679191 | 2664647 | 2678348 | 2680352 | 2674023 | 2676580 | 2684232 | 2683957 | 2680320 | 2675374 | 2680778 | 2679581 | 2678451 | 2680169 | 2676518 | 2678209 | 2679702 | 6 | 2683833.5821 | 8076.1477 | 0.0030092
7 | 3161898 | 3149645 | 3132690 | 3097153 | 3021705 | 3078035 | 3084394 | 3099096 | 3058990 | 3085354 | 3114015 | 3133721 | 3141053 | 3137151 | 3128253 | 3142478 | 3126533 | 3163724 | 3190869 | 3203831 | 3203831 | 3222093 | 3207930 | 3173484 | 3111008 | 3137594 | 3137689 | 3090808 | 3116789 | 3148835 | 3164327 | 3146853 | 3136026 | 3130227 | 3115222 | 3125220 | 3126673 | 3094688 | 3134956 | 3140423 | 3171366 | 3181570 | 3182541 | 3178635 | 3171121 | 3190294 | 3185186 | 3167322 | 3148399 | 3138325 | 3136685 | 3106938 | 3115086 | 3117826 | 3123971 | 3135387 | 3133890 | 3122215 | 3123859 | 3113601 | 3128990 | 3118361 | 3102698 | 3094030 | 3054215 | 3064640 | 3054751 | 7 | 3132495.0149 | 39391.9457 | 0.012575
8 | 3708036 | 3704370 | 3702826 | 3701552 | 3645464 | 3697339 | 3706512 | 3704799 | 3664442 | 3698068 | 3704755 | 3703832 | 3703575 | 3702343 | 3699190 | 3699067 | 3670540 | 3695660 | 3701581 | 3705840 | 3705840 | 3703294 | 3703075 | 3686360 | 3674271 | 3701457 | 3710747 | 3706219 | 3702143 | 3705296 | 3708159 | 3708447 | 3703599 | 3701341 | 3685822 | 3697512 | 3700926 | 3697541 | 3696970 | 3696827 | 3702130 | 3704030 | 3700683 | 3701453 | 3700066 | 3703979 | 3700322 | 3694816 | 3699864 | 3698910 | 3697818 | 3670428 | 3689743 | 3697955 | 3689748 | 3693763 | 3700141 | 3702381 | 3698646 | 3697288 | 3697950 | 3696334 | 3698451 | 3696298 | 3696803 | 3696326 | 3698278 | 8 | 3697615.5373 | 10762.7751 | 0.0029107
9 | 3315761 | 3311699 | 3313913 | 3309370 | 3260219 | 3311225 | 3308715 | 3315171 | 3282759 | 3310250 | 3315243 | 3313852 | 3313128 | 3310779 | 3306959 | 3307244 | 3281561 | 3304313 | 3311988 | 3315947 | 3315947 | 3312324 | 3308015 | 3297099 | 3288366 | 3312049 | 3308416 | 3310313 | 3307530 | 3306038 | 3306896 | 3308318 | 3307050 | 3307802 | 3289997 | 3305393 | 3309195 | 3308367 | 3303424 | 3305865 | 3304944 | 3309137 | 3309436 | 3308118 | 3308146 | 3306135 | 3302318 | 3305244 | 3305723 | 3302669 | 3302537 | 3284078 | 3298095 | 3306134 | 3298326 | 3304388 | 3307323 | 3303378 | 3302734 | 3305623 | 3302518 | 3304248 | 3305892 | 3305948 | 3301729 | 3306330 | 3305945 | 9 | 3305158.1493 | 9058.6644 | 0.0027408
10 | 3503350 | 3500557 | 3499935 | 3492747 | 3443065 | 3496274 | 3499951 | 3500788 | 3465282 | 3494932 | 3500360 | 3500874 | 3500705 | 3497356 | 3494668 | 3494680 | 3469696 | 3490352 | 3499817 | 3497835 | 3497835 | 3495821 | 3496625 | 3483240 | 3476749 | 3501703 | 3502333 | 3503238 | 3501021 | 3499401 | 3501000 | 3503570 | 3501743 | 3501559 | 3486785 | 3494876 | 3499834 | 3500024 | 3487030 | 3498237 | 3494636 | 3497482 | 3496660 | 3496613 | 3494518 | 3497927 | 3496376 | 3489659 | 3489498 | 3492997 | 3487388 | 3467189 | 3488462 | 3493778 | 3484794 | 3493679 | 3502031 | 3497211 | 3493741 | 3494281 | 3494982 | 3493346 | 3493589 | 3489992 | 3495170 | 3491504 | 3495585 | 10 | 3493715.4627 | 10009.4773 | 0.002865
11 | 3578103 | 3579795 | 3580166 | 3573078 | 3518008 | 3573105 | 3578916 | 3579840 | 3544271 | 3568979 | 3580773 | 3577983 | 3573953 | 3577669 | 3573152 | 3573720 | 3544874 | 3572396 | 3579300 | 3575131 | 3575131 | 3575080 | 3571894 | 3563763 | 3552756 | 3586221 | 3583848 | 3581912 | 3577792 | 3581784 | 3580223 | 3582366 | 3584484 | 3579997 | 3560933 | 3579121 | 3576745 | 3577275 | 3570143 | 3578536 | 3575193 | 3575850 | 3575069 | 3574916 | 3570851 | 3577008 | 3575996 | 3570351 | 3574693 | 3572761 | 3574534 | 3547704 | 3567850 | 3578376 | 3566408 | 3574524 | 3577195 | 3576486 | 3574803 | 3572707 | 3575789 | 3573712 | 3576552 | 3577309 | 3573889 | 3575673 | 3580218 | 11 | 3573278.1045 | 10579.2573 | 0.0029607
12 | 3208990 | 3227052 | 3345826 | 3319896 | 3222645 | 3276878 | 3286254 | 3287489 | 3262665 | 3287045 | 3274714 | 3271886 | 3277492 | 3282695 | 3288106 | 3278242 | 3237232 | 3267485 | 3286445 | 3282875 | 3282875 | 3291848 | 3281843 | 3269436 | 3256527 | 3284817 | 3274878 | 3277142 | 3275600 | 3279430 | 3276801 | 3280116 | 3290147 | 3283553 | 3263788 | 3271122 | 3272502 | 3262055 | 3266308 | 3281109 | 3282360 | 3281284 | 3271989 | 3277967 | 3268651 | 3279048 | 3278572 | 3276528 | 3275142 | 3274918 | 3278370 | 3265337 | 3269017 | 3272510 | 3276908 | 3280275 | 3278572 | 3277197 | 3269506 | 3279015 | 3278488 | 3285518 | 3289183 | 3294633 | 3284879 | 3279002 | 3283378 | 12 | 3276150.0896 | 18046.9687 | 0.0055086
13 | 3513454 | 3508768 | 3506057 | 3502298 | 3448196 | 3504697 | 3504839 | 3509867 | 3470896 | 3500768 | 3506063 | 3511520 | 3503093 | 3501582 | 3500007 | 3501780 | 3474802 | 3501311 | 3506655 | 3503898 | 3503898 | 3501806 | 3501098 | 3486563 | 3478883 | 3502313 | 3498877 | 3506094 | 3504624 | 3501934 | 3508070 | 3510152 | 3511038 | 3505877 | 3486191 | 3498465 | 3502769 | 3500181 | 3489003 | 3497524 | 3498592 | 3496547 | 3496208 | 3498414 | 3492722 | 3497212 | 3493957 | 3492324 | 3496133 | 3493304 | 3493339 | 3471302 | 3487794 | 3494256 | 3485057 | 3495051 | 3498567 | 3495253 | 3493529 | 3494277 | 3498574 | 3494620 | 3496994 | 3491012 | 3494788 | 3494036 | 3494460 | 13 | 3497078.1045 | 10585.912 | 0.0030271
14 | 3292209 | 3289396 | 3286757 | 3283456 | 3235660 | 3286193 | 3289253 | 3294866 | 3262082 | 3290059 | 3298737 | 3299188 | 3299710 | 3298982 | 3298813 | 3298253 | 3273040 | 3293249 | 3305202 | 3296438 | 3296438 | 3299121 | 3297273 | 3284879 | 3275010 | 3297949 | 3297856 | 3299454 | 3296290 | 3301092 | 3304746 | 3307497 | 3309975 | 3304069 | 3284727 | 3298370 | 3303360 | 3299088 | 3292431 | 3297619 | 3302131 | 3304619 | 3300002 | 3305005 | 3296611 | 3298394 | 3295087 | 3297063 | 3294890 | 3299490 | 3296624 | 3278252 | 3295546 | 3301017 | 3289722 | 3296909 | 3303423 | 3300745 | 3298617 | 3299482 | 3301862 | 3296410 | 3299775 | 3300331 | 3295671 | 3299249 | 3302540 | 14 | 3295048.5672 | 10936.3891 | 0.003319
15 | 3647841 | 3644519 | 3643456 | 3639397 | 3586463 | 3640252 | 3644234 | 3644392 | 3605801 | 3638809 | 3646373 | 3646904 | 3639500 | 3640712 | 3640853 | 3640378 | 3611989 | 3635633 | 3641650 | 3640145 | 3640145 | 3640340 | 3641025 | 3626375 | 3616406 | 3641467 | 3646309 | 3645272 | 3643736 | 3646754 | 3648702 | 3648206 | 3647721 | 3646490 | 3629953 | 3639896 | 3644107 | 3638994 | 3633232 | 3639050 | 3636565 | 3644149 | 3636261 | 3638108 | 3638057 | 3642458 | 3640291 | 3636024 | 3636749 | 3636196 | 3637042 | 3615914 | 3631171 | 3636899 | 3633400 | 3638212 | 3641475 | 3640855 | 3637879 | 3637797 | 3640979 | 3638549 | 3639023 | 3637444 | 3639544 | 3639219 | 3644352 | 15 | 3637881.9851 | 10156.0469 | 0.0027917
16 | 2817395 | 2808915 | 2817868 | 2835629 | 2773358 | 2863166 | 2850938 | 2858587 | 2804376 | 2930139 | 2885795 | 2857609 | 2885851 | 2917943 | 2804611 | 2808486 | 2821015 | 2820854 | 2847395 | 2831474 | 2831474 | 2815552 | 2832941 | 2807987 | 2730723 | 2708732 | 2719041 | 2744568 | 2748951 | 2748180 | 2695510 | 2716887 | 2673870 | 2675450 | 2668609 | 2707266 | 2700638 | 2661403 | 2641433 | 2645235 | 2629978 | 2632328 | 2636877 | 2637701 | 2609366 | 2628367 | 2671424 | 2660259 | 2644323 | 2656054 | 2658503 | 2655221 | 2691785 | 2709978 | 2735907 | 2721610 | 2723990 | 2740877 | 2786226 | 2792787 | 2814417 | 2779726 | 2782549 | 2734230 | 2747444 | 2704993 | 2717545 | 16 | 2749557 | 81060.3909 | 0.029481
17 | 3758349 | 3758664 | 3759846 | 3748724 | 3692912 | 3750102 | 3754768 | 3755214 | 3715592 | 3750992 | 3758055 | 3752270 | 3755171 | 3754122 | 3751464 | 3753945 | 3720295 | 3746348 | 3758083 | 3751813 | 3751813 | 3753552 | 3748879 | 3734833 | 3722826 | 3755370 | 3754773 | 3754196 | 3750643 | 3753030 | 3757983 | 3758988 | 3756637 | 3752358 | 3733286 | 3750772 | 3754256 | 3754242 | 3744670 | 3753389 | 3753409 | 3749464 | 3750189 | 3750349 | 3748520 | 3752499 | 3748414 | 3745381 | 3747836 | 3748120 | 3746206 | 3725792 | 3741764 | 3749545 | 3740141 | 3746459 | 3749620 | 3750514 | 3750538 | 3747797 | 3750999 | 3748325 | 3752774 | 3748855 | 3749688 | 3747807 | 3753018 | 17 | 3748317.1343 | 11009.4567 | 0.0029372
18 | 3960328 | 3954817 | 3955782 | 3949412 | 3892707 | 3952270 | 3956009 | 3956877 | 3916054 | 3949988 | 3954445 | 3958587 | 3953049 | 3951053 | 3949740 | 3952057 | 3920194 | 3943627 | 3953972 | 3954994 | 3954994 | 3951326 | 3950350 | 3936985 | 3921703 | 3953850 | 3955275 | 3958245 | 3955568 | 3957120 | 3954519 | 3959122 | 3956918 | 3950567 | 3935270 | 3951531 | 3953925 | 3955098 | 3946556 | 3951414 | 3951609 | 3949340 | 3952190 | 3949364 | 3948604 | 3951123 | 3949682 | 3949315 | 3946401 | 3945735 | 3941426 | 3921445 | 3941525 | 3947803 | 3938856 | 3949325 | 3950457 | 3948505 | 3947663 | 3949230 | 3953853 | 3943369 | 3947337 | 3948421 | 3945531 | 3947617 | 3952487 | 18 | 3947977.7761 | 11162.7602 | 0.0028275
19 | 3812856 | 3813336 | 3816907 | 3807430 | 3752377 | 3809797 | 3813142 | 3811968 | 3772907 | 3804571 | 3815741 | 3811014 | 3809789 | 3810263 | 3809732 | 3808921 | 3777323 | 3802766 | 3814047 | 3809096 | 3809096 | 3807858 | 3812528 | 3799337 | 3785124 | 3811864 | 3811667 | 3812448 | 3810016 | 3812831 | 3810992 | 3816821 | 3813676 | 3811900 | 3792085 | 3806163 | 3812574 | 3806789 | 3803744 | 3811222 | 3811485 | 3811713 | 3807314 | 3807133 | 3807050 | 3809449 | 3806186 | 3805672 | 3800522 | 3802108 | 3804000 | 3780604 | 3798510 | 3806465 | 3797119 | 3803894 | 3808067 | 3807404 | 3806612 | 3809656 | 3806538 | 3801419 | 3805481 | 3805046 | 3808211 | 3803327 | 3808324 | 19 | 3805552.6418 | 10614.7386 | 0.0027893
20 | 3651269 | 3650265 | 3651798 | 3638751 | 3585611 | 3644658 | 3650352 | 3648672 | 3611008 | 3643252 | 3647763 | 3650186 | 3648443 | 3652174 | 3645336 | 3645189 | 3619160 | 3639221 | 3651072 | 3646810 | 3646810 | 3645631 | 3644727 | 3629758 | 3621842 | 3642292 | 3649878 | 3650709 | 3647366 | 3649552 | 3650246 | 3651729 | 3650944 | 3649661 | 3629809 | 3647756 | 3650267 | 3645010 | 3635841 | 3645785 | 3642194 | 3647910 | 3645646 | 3644368 | 3639701 | 3643758 | 3643373 | 3642769 | 3638197 | 3644614 | 3642620 | 3621088 | 3638263 | 3643102 | 3636768 | 3642932 | 3648766 | 3643545 | 3643869 | 3647955 | 3644859 | 3642180 | 3643271 | 3641854 | 3640932 | 3638118 | 3644574 | 20 | 3642445.209 | 10605.1747 | 0.0029116
21 | 3554559 | 3552324 | 3551876 | 3543352 | 3495712 | 3547434 | 3547231 | 3552965 | 3516873 | 3543277 | 3550129 | 3552406 | 3549234 | 3548069 | 3548537 | 3546985 | 3519060 | 3543110 | 3546757 | 3548000 | 3548000 | 3552834 | 3546347 | 3531219 | 3522503 | 3550682 | 3547882 | 3548837 | 3547273 | 3548925 | 3551439 | 3551258 | 3550777 | 3545380 | 3530810 | 3544014 | 3548762 | 3548488 | 3543082 | 3547168 | 3546238 | 3548075 | 3542123 | 3546827 | 3542832 | 3548648 | 3543128 | 3538969 | 3538445 | 3539188 | 3538959 | 3518288 | 3533676 | 3542354 | 3528395 | 3541333 | 3545721 | 3542047 | 3540497 | 3544443 | 3539665 | 3538213 | 3539732 | 3543841 | 3540929 | 3541243 | 3547444 | 21 | 3542758.5522 | 9975.1073 | 0.0028156
22 | 3124984 | 3126626 | 3123712 | 3121561 | 3082278 | 3124198 | 3125545 | 3127632 | 3095219 | 3121330 | 3121423 | 3126995 | 3121646 | 3124699 | 3122481 | 3119882 | 3099353 | 3117402 | 3123672 | 3123660 | 3123660 | 3121743 | 3120890 | 3109600 | 3098194 | 3123709 | 3123760 | 3120398 | 3119925 | 3122930 | 3118860 | 3127201 | 3127143 | 3121645 | 3107993 | 3123554 | 3126103 | 3116684 | 3115967 | 3118126 | 3119538 | 3120308 | 3118829 | 3116603 | 3113251 | 3114729 | 3115370 | 3112647 | 3111400 | 3111397 | 3113134 | 3092576 | 3110520 | 3114839 | 3107918 | 3111789 | 3117970 | 3114660 | 3115259 | 3114168 | 3113697 | 3112670 | 3113832 | 3113949 | 3112181 | 3116282 | 3115170 | 22 | 3116821.9254 | 8516.3868 | 0.0027324
23 | 3544084 | 3543950 | 3537424 | 3533503 | 3485091 | 3534182 | 3544653 | 3541441 | 3509450 | 3535553 | 3541572 | 3543678 | 3537161 | 3538647 | 3534320 | 3535044 | 3507840 | 3532564 | 3538922 | 3536091 | 3536091 | 3534606 | 3534411 | 3524505 | 3515327 | 3538570 | 3540167 | 3542010 | 3536405 | 3532491 | 3539804 | 3538324 | 3540690 | 3539713 | 3519952 | 3531366 | 3537338 | 3539223 | 3531184 | 3536073 | 3536234 | 3533368 | 3531228 | 3532791 | 3530868 | 3529019 | 3528006 | 3527607 | 3531915 | 3527490 | 3527873 | 3509421 | 3526638 | 3529233 | 3523251 | 3528244 | 3534447 | 3533910 | 3529938 | 3531079 | 3533201 | 3531148 | 3532027 | 3533156 | 3529475 | 3531975 | 3536337 | 23 | 3532138.791 | 9552.3973 | 0.0027044
24 | 3618324 | 3613179 | 3614876 | 3607360 | 3556533 | 3611774 | 3617171 | 3613640 | 3578438 | 3607783 | 3615163 | 3612010 | 3611616 | 3611015 | 3611122 | 3613886 | 3584441 | 3604249 | 3609088 | 3610518 | 3610518 | 3608530 | 3608962 | 3595742 | 3584487 | 3608295 | 3614258 | 3615667 | 3612847 | 3609326 | 3610015 | 3611334 | 3612716 | 3607332 | 3596051 | 3609400 | 3612387 | 3609467 | 3605121 | 3610657 | 3606370 | 3610613 | 3606322 | 3609622 | 3606212 | 3611795 | 3608355 | 3609532 | 3605296 | 3604402 | 3605279 | 3586468 | 3600095 | 3607211 | 3599101 | 3604115 | 3611656 | 3606504 | 3605264 | 3607251 | 3608508 | 3602173 | 3611401 | 3605801 | 3607656 | 3606822 | 3606531 | 24 | 3606502.2836 | 9626.8112 | 0.0026693
25 | 3856103 | 3850801 | 3850222 | 3847366 | 3787550 | 3847055 | 3850490 | 3851241 | 3810912 | 3845041 | 3847914 | 3852276 | 3847423 | 3845590 | 3845458 | 3848345 | 3815215 | 3842662 | 3850650 | 3851883 | 3851883 | 3847439 | 3848421 | 3830114 | 3816436 | 3848270 | 3846374 | 3852856 | 3850076 | 3845992 | 3852200 | 3851713 | 3850250 | 3849049 | 3832805 | 3845897 | 3847091 | 3845458 | 3839524 | 3847622 | 3846992 | 3845311 | 3846377 | 3844494 | 3842832 | 3845805 | 3846308 | 3836825 | 3840203 | 3839669 | 3841375 | 3823935 | 3836283 | 3842343 | 3833688 | 3839631 | 3846727 | 3842923 | 3844144 | 3845038 | 3843488 | 3842736 | 3842924 | 3844080 | 3844809 | 3844854 | 3845486 | 25 | 3843029.0597 | 10859.2293 | 0.0028257
26 | 3989888 | 3984114 | 3984766 | 3976382 | 3921600 | 3978087 | 3986886 | 3985709 | 3948437 | 3981711 | 3984919 | 3988016 | 3982665 | 3980341 | 3981022 | 3981370 | 3950452 | 3976171 | 3982832 | 3986255 | 3986255 | 3984738 | 3983653 | 3966229 | 3954500 | 3983385 | 3982750 | 3986634 | 3981199 | 3985161 | 3984982 | 3985239 | 3982312 | 3983575 | 3965412 | 3977804 | 3985682 | 3984272 | 3972405 | 3980444 | 3982482 | 3984825 | 3977702 | 3979881 | 3975851 | 3980564 | 3977845 | 3975412 | 3975888 | 3975871 | 3969499 | 3951116 | 3973380 | 3976263 | 3966497 | 3973644 | 3977910 | 3979315 | 3974450 | 3976159 | 3979070 | 3976125 | 3977091 | 3976225 | 3974335 | 3980396 | 3979128 | 26 | 3977241.4627 | 10946.868 | 0.0027524
27 | 3782370 | 3784397 | 3781528 | 3776749 | 3722667 | 3775893 | 3781277 | 3783544 | 3744527 | 3774633 | 3783210 | 3777982 | 3777068 | 3782785 | 3779704 | 3780037 | 3746817 | 3774585 | 3779287 | 3780602 | 3780602 | 3779916 | 3774102 | 3761794 | 3754944 | 3777850 | 3778986 | 3782833 | 3777839 | 3777711 | 3781088 | 3783110 | 3783152 | 3781367 | 3761638 | 3772881 | 3780036 | 3778812 | 3766599 | 3776209 | 3776848 | 3776364 | 3773684 | 3772765 | 3771588 | 3777741 | 3773983 | 3771267 | 3772382 | 3772023 | 3771070 | 3750592 | 3771831 | 3774526 | 3762204 | 3772261 | 3776481 | 3770792 | 3772262 | 3773898 | 3775682 | 3772774 | 3773727 | 3772461 | 3769626 | 3773813 | 3773782 | 27 | 3773545.6418 | 10241.2216 | 0.002714
28 | 3911195 | 3910379 | 3909368 | 3902524 | 3848623 | 3905793 | 3907440 | 3910887 | 3874648 | 3901987 | 3911448 | 3909022 | 3909395 | 3908925 | 3903222 | 3909036 | 3876666 | 3899213 | 3910427 | 3904860 | 3904860 | 3907052 | 3901615 | 3890513 | 3878103 | 3907458 | 3908431 | 3910662 | 3906936 | 3905929 | 3911926 | 3911229 | 3906218 | 3910809 | 3890816 | 3905571 | 3907765 | 3905190 | 3898442 | 3902463 | 3907036 | 3904894 | 3899611 | 3901552 | 3902651 | 3905627 | 3902407 | 3901496 | 3902280 | 3899868 | 3898907 | 3873720 | 3896286 | 3900542 | 3895880 | 3902070 | 3906460 | 3904797 | 3902697 | 3898879 | 3902171 | 3900982 | 3901841 | 3900315 | 3905570 | 3900877 | 3903323 | 28 | 3901787.8358 | 10467.4293 | 0.0026827
29 | 3925597 | 3923672 | 3925018 | 3916536 | 3863146 | 3919341 | 3926046 | 3926805 | 3887351 | 3918591 | 3925088 | 3926934 | 3924984 | 3924172 | 3922229 | 3923164 | 3890377 | 3911886 | 3922615 | 3921089 | 3921089 | 3919115 | 3922791 | 3910534 | 3894767 | 3921781 | 3924400 | 3921555 | 3923127 | 3924729 | 3925204 | 3926828 | 3925650 | 3926232 | 3903591 | 3917264 | 3925649 | 3919651 | 3911923 | 3918962 | 3919832 | 3922019 | 3917134 | 3919043 | 3919891 | 3918127 | 3919918 | 3915728 | 3916241 | 3913616 | 3912506 | 3890708 | 3912251 | 3916364 | 3907906 | 3913259 | 3916937 | 3918871 | 3917262 | 3921641 | 3914334 | 3914723 | 3916241 | 3911537 | 3915029 | 3916294 | 3920499 | 29 | 3916826.7761 | 10692.8231 | 0.00273
30 | 4231224 | 4234445 | 4233738 | 4223593 | 4165045 | 4230707 | 4241095 | 4234490 | 4190860 | 4229169 | 4235672 | 4236513 | 4232137 | 4231194 | 4230316 | 4232065 | 4192831 | 4225030 | 4230982 | 4229957 | 4229957 | 4231003 | 4228221 | 4209742 | 4198358 | 4230101 | 4236130 | 4234126 | 4231494 | 4233585 | 4238378 | 4235695 | 4233693 | 4234534 | 4214489 | 4228462 | 4235243 | 4231975 | 4224232 | 4227519 | 4228069 | 4230112 | 4225492 | 4227842 | 4226929 | 4226153 | 4223899 | 4222369 | 4219129 | 4224202 | 4223723 | 4197663 | 4218426 | 4224398 | 4212289 | 4224227 | 4228664 | 4228465 | 4224633 | 4222989 | 4227818 | 4222937 | 4224331 | 4223527 | 4222262 | 4228294 | 4227881 | 30 | 4225294.3731 | 12273.1021 | 0.0029047
31 | 4082495 | 4081868 | 4079509 | 4070221 | 4013059 | 4077630 | 4080034 | 4079853 | 4038663 | 4073067 | 4079272 | 4080287 | 4076835 | 4077775 | 4074830 | 4074599 | 4042312 | 4069204 | 4074406 | 4076002 | 4076002 | 4076117 | 4074472 | 4061357 | 4049158 | 4078828 | 4079245 | 4076822 | 4075796 | 4078855 | 4076710 | 4077190 | 4078934 | 4079063 | 4056518 | 4074513 | 4077702 | 4075937 | 4067969 | 4072071 | 4072783 | 4074241 | 4074412 | 4073475 | 4069907 | 4076172 | 4071870 | 4070094 | 4068309 | 4069012 | 4070318 | 4048232 | 4067945 | 4074169 | 4063202 | 4067992 | 4073157 | 4071430 | 4072979 | 4069520 | 4071715 | 4068906 | 4073545 | 4068693 | 4070439 | 4075609 | 4074594 | 31 | 4071162.6866 | 11050.5928 | 0.0027144
32 | 4056077 | 4046893 | 4045367 | 4038384 | 3984811 | 4041171 | 4048284 | 4045051 | 4004977 | 4035907 | 4048478 | 4050799 | 4046580 | 4048171 | 4042401 | 4045449 | 4008216 | 4036593 | 4048646 | 4040048 | 4040048 | 4044352 | 4040499 | 4027544 | 4013382 | 4042245 | 4040849 | 4045096 | 4040336 | 4042634 | 4040840 | 4047511 | 4043818 | 4041132 | 4026055 | 4038899 | 4044043 | 4038768 | 4030346 | 4038826 | 4039018 | 4039261 | 4042226 | 4042306 | 4038834 | 4041016 | 4038115 | 4036726 | 4032152 | 4032568 | 4031998 | 4012103 | 4029912 | 4038584 | 4027580 | 4037597 | 4040979 | 4040527 | 4034591 | 4036445 | 4036469 | 4038627 | 4038044 | 4037693 | 4032499 | 4040108 | 4042750 | 32 | 4037436.6269 | 11195.0202 | 0.0027728
33 | 3460139 | 3462475 | 3459926 | 3459357 | 3410983 | 3460709 | 3464161 | 3464042 | 3429309 | 3459823 | 3463528 | 3459467 | 3458098 | 3458640 | 3458493 | 3456629 | 3432356 | 3452867 | 3461013 | 3458387 | 3458387 | 3459702 | 3456819 | 3448194 | 3437342 | 3457042 | 3458388 | 3459684 | 3454659 | 3457426 | 3455864 | 3459285 | 3458397 | 3460388 | 3445783 | 3458701 | 3467253 | 3463133 | 3452471 | 3456238 | 3457700 | 3457450 | 3451254 | 3453935 | 3450051 | 3453670 | 3450526 | 3451186 | 3446804 | 3449940 | 3447186 | 3425712 | 3443147 | 3449781 | 3441627 | 3447754 | 3451753 | 3451182 | 3448418 | 3450210 | 3451343 | 3450364 | 3453461 | 3446309 | 3451308 | 3453305 | 3455532 | 33 | 3453081.5821 | 9412.1149 | 0.0027257
34 | 3833075 | 3829720 | 3834263 | 3825142 | 3775872 | 3828850 | 3834010 | 3833484 | 3794982 | 3827331 | 3831280 | 3835690 | 3829886 | 3832815 | 3829542 | 3830775 | 3800247 | 3824624 | 3834543 | 3826842 | 3826842 | 3826607 | 3829857 | 3815045 | 3801320 | 3828769 | 3831694 | 3832202 | 3828947 | 3829196 | 3830384 | 3833295 | 3831849 | 3829025 | 3812417 | 3827066 | 3830711 | 3827367 | 3822251 | 3825872 | 3824848 | 3828469 | 3824912 | 3825306 | 3818633 | 3827382 | 3820703 | 3821600 | 3818924 | 3823063 | 3817974 | 3795840 | 3813673 | 3819870 | 3814114 | 3818145 | 3824453 | 3821010 | 3820176 | 3822784 | 3824528 | 3818382 | 3820475 | 3821268 | 3818775 | 3821643 | 3824862 | 34 | 3823365.6866 | 10371.8876 | 0.0027128
35 | 4313321 | 4311743 | 4306580 | 4300882 | 4242493 | 4306133 | 4311944 | 4309742 | 4269921 | 4302448 | 4307291 | 4308385 | 4307247 | 4308759 | 4305580 | 4304704 | 4271806 | 4296670 | 4307885 | 4307888 | 4307888 | 4308004 | 4308066 | 4288871 | 4277662 | 4305893 | 4311725 | 4309632 | 4306260 | 4307062 | 4309620 | 4314404 | 4309303 | 4310291 | 4285915 | 4304403 | 4310532 | 4304874 | 4298261 | 4307790 | 4308171 | 4304561 | 4303498 | 4302521 | 4300285 | 4305864 | 4301754 | 4299550 | 4297725 | 4299915 | 4297751 | 4270236 | 4289685 | 4303366 | 4289042 | 4297026 | 4302962 | 4303948 | 4298428 | 4297177 | 4301197 | 4297641 | 4303002 | 4300558 | 4295862 | 4301102 | 4302650 | 35 | 4300945.597 | 11888.0797 | 0.0027641
36 | 3285032 | 3290040 | 3289056 | 3280752 | 3244890 | 3294917 | 3300337 | 3296284 | 3265571 | 3289006 | 3299848 | 3302544 | 3291921 | 3277386 | 3281784 | 3279296 | 3249989 | 3271710 | 3276237 | 3275902 | 3275902 | 3273692 | 3270118 | 3262775 | 3248460 | 3269318 | 3273644 | 3277904 | 3282772 | 3286078 | 3288217 | 3292370 | 3294053 | 3287010 | 3270673 | 3281995 | 3284314 | 3279675 | 3261805 | 3258815 | 3262163 | 3261914 | 3256739 | 3258778 | 3257681 | 3260832 | 3257678 | 3258792 | 3262884 | 3266082 | 3264299 | 3243831 | 3253192 | 3267605 | 3255337 | 3254479 | 3260865 | 3265735 | 3257977 | 3262635 | 3262497 | 3261925 | 3255281 | 3255889 | 3257389 | 3249877 | 3252532 | 36 | 3270850.4478 | 14993.8159 | 0.0045841
37 | 3100433 | 3093435 | 3089288 | 3085746 | 3038955 | 3080909 | 3075498 | 3080077 | 3060420 | 3085778 | 3086305 | 3079147 | 3080736 | 3082225 | 3078059 | 3078200 | 3059475 | 3074771 | 3080461 | 3076590 | 3076590 | 3087791 | 3084729 | 3068317 | 3045325 | 3065761 | 3052304 | 3056593 | 3049092 | 3058480 | 3067649 | 3050030 | 3046793 | 3042250 | 3022572 | 3025876 | 3024046 | 3024072 | 3013195 | 3020180 | 3020841 | 3012829 | 3018748 | 3018750 | 3016979 | 3008669 | 3012556 | 3013421 | 3010418 | 3014071 | 3009998 | 2994968 | 3011454 | 3017775 | 3014965 | 3009796 | 3005772 | 3019640 | 3011279 | 3008535 | 3010894 | 3006551 | 3000612 | 2996656 | 3003645 | 2994840 | 3000591 | 37 | 3041991.1343 | 32191.1095 | 0.010582
38 | 4318335 | 4319274 | 4321749 | 4313632 | 4247526 | 4313747 | 4323755 | 4316905 | 4274855 | 4313210 | 4320435 | 4320425 | 4319240 | 4315527 | 4310237 | 4312893 | 4277442 | 4308594 | 4318233 | 4317725 | 4317725 | 4314343 | 4311800 | 4300720 | 4286384 | 4316989 | 4317797 | 4319209 | 4314674 | 4317467 | 4317810 | 4320772 | 4318524 | 4314608 | 4298035 | 4315800 | 4321243 | 4317322 | 4304432 | 4315384 | 4311056 | 4313182 | 4312181 | 4309823 | 4305470 | 4314723 | 4310466 | 4307180 | 4307603 | 4308099 | 4306765 | 4282629 | 4302628 | 4305913 | 4298701 | 4308447 | 4316242 | 4311554 | 4311732 | 4306405 | 4309129 | 4307089 | 4307532 | 4310727 | 4309611 | 4308875 | 4311435 | 38 | 4309850.3582 | 12238.3997 | 0.0028396
39 | 4572637 | 4567924 | 4562952 | 4559334 | 4500941 | 4568425 | 4577698 | 4573963 | 4528282 | 4566996 | 4570674 | 4573109 | 4570419 | 4575628 | 4569811 | 4572537 | 4535719 | 4560852 | 4576664 | 4571839 | 4571839 | 4574042 | 4571203 | 4555313 | 4540546 | 4573063 | 4573451 | 4574648 | 4576284 | 4576190 | 4578733 | 4578285 | 4579605 | 4572637 | 4554325 | 4569782 | 4577130 | 4571715 | 4565166 | 4571559 | 4574390 | 4571029 | 4567399 | 4572386 | 4569468 | 4574573 | 4569228 | 4567135 | 4562834 | 4566364 | 4568671 | 4544478 | 4565425 | 4572473 | 4562354 | 4567646 | 4572248 | 4572408 | 4569461 | 4572254 | 4569073 | 4568678 | 4571588 | 4568857 | 4569572 | 4567306 | 4569254 | 39 | 4567290.6269 | 12417.3381 | 0.0027188
40 | 4768931 | 4768815 | 4770633 | 4759085 | 4698730 | 4767198 | 4770688 | 4769470 | 4722398 | 4760779 | 4770878 | 4772457 | 4766451 | 4769938 | 4764468 | 4767478 | 4729395 | 4761404 | 4768845 | 4765591 | 4765591 | 4767358 | 4767823 | 4749803 | 4729792 | 4766015 | 4767788 | 4770348 | 4765310 | 4766931 | 4767812 | 4772308 | 4766049 | 4762490 | 4745452 | 4763564 | 4770346 | 4761059 | 4754449 | 4763879 | 4764723 | 4765175 | 4762548 | 4763309 | 4758421 | 4762189 | 4766675 | 4753893 | 4756587 | 4758080 | 4754332 | 4733200 | 4751233 | 4759282 | 4747428 | 4757616 | 4763156 | 4757309 | 4758765 | 4755629 | 4763950 | 4754271 | 4758350 | 4755799 | 4758907 | 4757877 | 4758075 | 40 | 4759620.5672 | 12565.8349 | 0.0026401
41 | 4572106 | 4568887 | 4572194 | 4555827 | 4493437 | 4560946 | 4569820 | 4568677 | 4523292 | 4560733 | 4568756 | 4568766 | 4565211 | 4564338 | 4563493 | 4561826 | 4528631 | 4558742 | 4565978 | 4572098 | 4572098 | 4564606 | 4561164 | 4546043 | 4534172 | 4566815 | 4564668 | 4571904 | 4564088 | 4569067 | 4564894 | 4571046 | 4567573 | 4566166 | 4541211 | 4559468 | 4568028 | 4565945 | 4556194 | 4564758 | 4565977 | 4558640 | 4556850 | 4562031 | 4556659 | 4561641 | 4559933 | 4555293 | 4557912 | 4556085 | 4555175 | 4533729 | 4553271 | 4557987 | 4549848 | 4557118 | 4562065 | 4561018 | 4559700 | 4556452 | 4559033 | 4558064 | 4559540 | 4559874 | 4555476 | 4559680 | 4562315 | 41 | 4559030.3284 | 12675.1306 | 0.0027802
42 | 4941533 | 4935124 | 4938352 | 4929052 | 4865322 | 4932277 | 4936574 | 4938567 | 4892288 | 4929259 | 4940125 | 4938006 | 4932966 | 4932662 | 4927905 | 4935761 | 4894508 | 4927477 | 4936547 | 4937534 | 4937534 | 4938304 | 4934208 | 4913179 | 4902790 | 4934541 | 4936286 | 4936722 | 4936562 | 4936183 | 4938370 | 4942724 | 4942173 | 4938176 | 4915292 | 4932469 | 4940462 | 4931806 | 4922615 | 4932220 | 4933005 | 4935523 | 4927510 | 4927206 | 4925651 | 4933743 | 4930158 | 4925562 | 4922125 | 4925654 | 4923005 | 4894539 | 4920320 | 4926018 | 4916594 | 4925309 | 4931305 | 4933021 | 4928016 | 4927496 | 4926857 | 4921008 | 4929543 | 4924824 | 4922719 | 4927620 | 4928172 | 42 | 4928014.2985 | 13089.6989 | 0.0026562
43 | 4887724 | 4884897 | 4886211 | 4878018 | 4804221 | 4869744 | 4882908 | 4879936 | 4835167 | 4874633 | 4883055 | 4884792 | 4878709 | 4880266 | 4875347 | 4877123 | 4836537 | 4870093 | 4881228 | 4882626 | 4882626 | 4881030 | 4881033 | 4866519 | 4854383 | 4890116 | 4892483 | 4898550 | 4893570 | 4894056 | 4899845 | 4900153 | 4897066 | 4898187 | 4875066 | 4885417 | 4895734 | 4893027 | 4881811 | 4885307 | 4890134 | 4892296 | 4879188 | 4881608 | 4877224 | 4880024 | 4878354 | 4873547 | 4879458 | 4877224 | 4879084 | 4853621 | 4877395 | 4886238 | 4870789 | 4880505 | 4884866 | 4882896 | 4877434 | 4882603 | 4885729 | 4879171 | 4884058 | 4877714 | 4878481 | 4880168 | 4880720 | 43 | 4879548.403 | 14997.7303 | 0.0030736
44 | 4906241 | 4905071 | 4908438 | 4892749 | 4832924 | 4897978 | 4901766 | 4904343 | 4861022 | 4893025 | 4904063 | 4904814 | 4899352 | 4904461 | 4896370 | 4903561 | 4863082 | 4890232 | 4904166 | 4902151 | 4902151 | 4899962 | 4899330 | 4883058 | 4863401 | 4900605 | 4902295 | 4901733 | 4898609 | 4895146 | 4901367 | 4903840 | 4904844 | 4900416 | 4875617 | 4897127 | 4903465 | 4896739 | 4889011 | 4898877 | 4899873 | 4902361 | 4899302 | 4895680 | 4891865 | 4895234 | 4892091 | 4891986 | 4890355 | 4889778 | 4888220 | 4857936 | 4886177 | 4887594 | 4881322 | 4887664 | 4893724 | 4895137 | 4888675 | 4891163 | 4891038 | 4890541 | 4888946 | 4895317 | 4889040 | 4889065 | 4890483 | 44 | 4893044.3134 | 12909.1877 | 0.0026383
45 | 4346750 | 4347681 | 4346315 | 4338654 | 4285388 | 4341953 | 4349938 | 4346152 | 4308346 | 4344232 | 4345379 | 4346850 | 4346238 | 4342152 | 4339264 | 4342705 | 4311574 | 4334668 | 4346399 | 4343233 | 4343233 | 4343344 | 4338672 | 4327339 | 4314979 | 4341220 | 4341288 | 4345958 | 4342848 | 4338985 | 4340394 | 4343251 | 4340657 | 4338909 | 4319459 | 4336817 | 4339435 | 4340080 | 4332484 | 4338293 | 4336214 | 4337948 | 4338380 | 4335108 | 4333232 | 4340717 | 4338700 | 4332881 | 4331558 | 4333725 | 4327826 | 4307885 | 4326272 | 4331128 | 4323094 | 4331267 | 4339071 | 4332352 | 4327903 | 4331512 | 4329536 | 4332993 | 4332615 | 4329121 | 4329042 | 4329987 | 4336161 | 45 | 4335190.209 | 10943.1438 | 0.0025243
46 | 3197659 | 3192673 | 3196560 | 3185358 | 3151594 | 3195548 | 3200297 | 3198870 | 3164806 | 3191178 | 3192654 | 3195010 | 3191123 | 3194141 | 3191523 | 3190339 | 3166210 | 3186331 | 3194365 | 3189337 | 3189337 | 3191815 | 3189931 | 3174970 | 3166640 | 3191693 | 3191226 | 3190454 | 3185802 | 3185936 | 3186799 | 3191358 | 3188681 | 3187545 | 3175151 | 3186779 | 3189825 | 3182838 | 3177604 | 3182634 | 3179234 | 3184167 | 3175311 | 3176630 | 3172682 | 3178139 | 3172903 | 3169043 | 3170671 | 3168271 | 3169717 | 3154779 | 3165686 | 3172927 | 3164691 | 3170149 | 3175301 | 3173755 | 3174767 | 3173181 | 3173324 | 3171097 | 3174711 | 3173392 | 3175609 | 3176672 | 3178289 | 46 | 3181159.5821 | 10920.2973 | 0.0034328
47 | 3627872 | 3626003 | 3624250 | 3620934 | 3567057 | 3623595 | 3625062 | 3621583 | 3590319 | 3623099 | 3626711 | 3625315 | 3622901 | 3623596 | 3620562 | 3623711 | 3590555 | 3616268 | 3627500 | 3623607 | 3623607 | 3626326 | 3621365 | 3609183 | 3597569 | 3621296 | 3627849 | 3626907 | 3624759 | 3623770 | 3626559 | 3626804 | 3623316 | 3622969 | 3604455 | 3621396 | 3624154 | 3620604 | 3615395 | 3620289 | 3620669 | 3620185 | 3620760 | 3619843 | 3617041 | 3620363 | 3619007 | 3620460 | 3621532 | 3618351 | 3617396 | 3596316 | 3613855 | 3618493 | 3608140 | 3617121 | 3623614 | 3623797 | 3619437 | 3620533 | 3617319 | 3617236 | 3618462 | 3620120 | 3617930 | 3619050 | 3620813 | 47 | 3618610.6716 | 10088.9401 | 0.0027881
48 | 3397230 | 3393541 | 3392043 | 3386325 | 3337786 | 3383768 | 3388488 | 3394380 | 3358249 | 3384894 | 3387970 | 3387504 | 3387059 | 3391979 | 3386376 | 3390440 | 3356490 | 3377857 | 3391033 | 3391448 | 3391448 | 3387194 | 3386557 | 3372980 | 3361154 | 3388725 | 3388856 | 3390121 | 3387641 | 3392310 | 3390592 | 3394512 | 3395682 | 3386297 | 3372272 | 3387487 | 3394129 | 3384547 | 3385508 | 3388967 | 3386793 | 3384887 | 3383828 | 3386985 | 3379607 | 3389500 | 3385631 | 3383168 | 3385369 | 3379467 | 3382155 | 3361895 | 3377402 | 3381948 | 3373397 | 3382651 | 3387181 | 3387283 | 3387227 | 3384104 | 3382767 | 3384492 | 3385565 | 3382984 | 3381914 | 3385766 | 3387623 | 48 | 3384081.0149 | 9903.6555 | 0.0029265
49 | 3375215 | 3375708 | 3374007 | 3372318 | 3318428 | 3369947 | 3376376 | 3375946 | 3341759 | 3369038 | 3377053 | 3377425 | 3374389 | 3372940 | 3369859 | 3375219 | 3344929 | 3369637 | 3374739 | 3374538 | 3374538 | 3374150 | 3369139 | 3356669 | 3347566 | 3369645 | 3377228 | 3380109 | 3371793 | 3374036 | 3376296 | 3380897 | 3380448 | 3373516 | 3355898 | 3369942 | 3377182 | 3371305 | 3366676 | 3371557 | 3370693 | 3374849 | 3370835 | 3372299 | 3366829 | 3370597 | 3369351 | 3370015 | 3370095 | 3367694 | 3367311 | 3347042 | 3364506 | 3367485 | 3361205 | 3368009 | 3374576 | 3369566 | 3364770 | 3370530 | 3369664 | 3365400 | 3373565 | 3370220 | 3366448 | 3366119 | 3369831 | 49 | 3368918.8657 | 9937.681 | 0.0029498
50 | 3506300 | 3501639 | 3505186 | 3497319 | 3443945 | 3501395 | 3504328 | 3506451 | 3469424 | 3500625 | 3504814 | 3504913 | 3503905 | 3503117 | 3500975 | 3501342 | 3471950 | 3494487 | 3504730 | 3500418 | 3500418 | 3503916 | 3501721 | 3490621 | 3474302 | 3501180 | 3503622 | 3505219 | 3504584 | 3506328 | 3505607 | 3506459 | 3504422 | 3502887 | 3483591 | 3498758 | 3504212 | 3498284 | 3492921 | 3498036 | 3501541 | 3499779 | 3498494 | 3499410 | 3497970 | 3502007 | 3501316 | 3498363 | 3497680 | 3497255 | 3495558 | 3477848 | 3490741 | 3500015 | 3491989 | 3500916 | 3502550 | 3501217 | 3495269 | 3499884 | 3497859 | 3496217 | 3500178 | 3496820 | 3497930 | 3497345 | 3499840 | 50 | 3497766.2985 | 10129.9685 | 0.0028961
51 | 3541153 | 3536350 | 3538614 | 3532731 | 3476965 | 3529515 | 3537547 | 3538728 | 3502794 | 3532607 | 3536702 | 3539769 | 3537540 | 3537844 | 3535580 | 3534362 | 3507543 | 3528334 | 3537197 | 3536432 | 3536432 | 3541220 | 3533813 | 3520853 | 3510898 | 3533740 | 3536057 | 3540451 | 3536556 | 3538707 | 3536390 | 3538348 | 3539415 | 3533523 | 3522300 | 3534376 | 3540044 | 3534432 | 3526804 | 3533702 | 3538257 | 3535094 | 3534885 | 3531648 | 3530459 | 3535320 | 3533430 | 3528149 | 3529095 | 3534979 | 3533097 | 3510559 | 3526359 | 3533202 | 3522163 | 3529008 | 3535794 | 3532160 | 3530504 | 3530719 | 3532494 | 3529264 | 3533241 | 3530939 | 3531254 | 3535624 | 3536520 | 51 | 3531501.2687 | 10102.0229 | 0.0028605
52 | 3727653 | 3731184 | 3729601 | 3723760 | 3665758 | 3723954 | 3729938 | 3730345 | 3691282 | 3720653 | 3730729 | 3728867 | 3721578 | 3720792 | 3722210 | 3723933 | 3693302 | 3716562 | 3728940 | 3725005 | 3725005 | 3727819 | 3724944 | 3710174 | 3698375 | 3727271 | 3728915 | 3726879 | 3724927 | 3723775 | 3725308 | 3730218 | 3728337 | 3724867 | 3705257 | 3718454 | 3725866 | 3723200 | 3719313 | 3726706 | 3722915 | 3723023 | 3723033 | 3724252 | 3716142 | 3722440 | 3719576 | 3718201 | 3717401 | 3718048 | 3713777 | 3694973 | 3716337 | 3717626 | 3712426 | 3711923 | 3723739 | 3722267 | 3718856 | 3723706 | 3721996 | 3716300 | 3720725 | 3719633 | 3717584 | 3724187 | 3721507 | 52 | 3719914.1642 | 10733.0189 | 0.0028853
53 | 3283854 | 3282187 | 3279459 | 3272906 | 3223325 | 3277240 | 3280098 | 3278093 | 3247736 | 3275473 | 3279253 | 3279948 | 3274837 | 3282534 | 3277121 | 3279678 | 3253084 | 3269998 | 3277828 | 3280030 | 3280030 | 3278458 | 3274496 | 3266186 | 3256062 | 3278805 | 3280175 | 3279058 | 3278359 | 3277155 | 3282472 | 3282227 | 3284062 | 3279511 | 3263305 | 3275161 | 3280204 | 3276446 | 3270135 | 3276721 | 3273892 | 3278345 | 3273589 | 3277729 | 3272900 | 3275843 | 3274606 | 3274323 | 3272786 | 3273959 | 3268283 | 3251529 | 3266353 | 3273637 | 3261310 | 3271147 | 3280615 | 3276573 | 3277380 | 3272451 | 3272141 | 3271493 | 3273504 | 3275406 | 3276864 | 3275624 | 3275500 | 53 | 3273753.6119 | 9542.7438 | 0.0029149
54 | 3262469 | 3264896 | 3263310 | 3257395 | 3210756 | 3261640 | 3264795 | 3264336 | 3229534 | 3253371 | 3261607 | 3260778 | 3261971 | 3262355 | 3259479 | 3263606 | 3232773 | 3255045 | 3263118 | 3258817 | 3258817 | 3259994 | 3260538 | 3248525 | 3239044 | 3262913 | 3263048 | 3264067 | 3261555 | 3262406 | 3261764 | 3269746 | 3263982 | 3259050 | 3242956 | 3253603 | 3261309 | 3260433 | 3252589 | 3256181 | 3258350 | 3258570 | 3260936 | 3260247 | 3256820 | 3261620 | 3257961 | 3258320 | 3253649 | 3254275 | 3257521 | 3236146 | 3254508 | 3255449 | 3250638 | 3258788 | 3261106 | 3257586 | 3257078 | 3258817 | 3264023 | 3259951 | 3254950 | 3256377 | 3252916 | 3260774 | 3260218 | 54 | 3256897.9851 | 9237.0051 | 0.0028361
55 | 3523702 | 3526412 | 3526489 | 3512794 | 3462983 | 3517893 | 3521505 | 3523944 | 3485545 | 3516509 | 3522484 | 3521771 | 3519598 | 3523467 | 3517939 | 3519882 | 3491373 | 3516557 | 3523109 | 3520515 | 3520515 | 3518080 | 3516398 | 3505597 | 3494358 | 3521827 | 3520056 | 3521475 | 3522431 | 3517371 | 3522511 | 3519291 | 3522301 | 3517941 | 3503561 | 3516676 | 3521958 | 3516965 | 3514262 | 3519182 | 3518635 | 3520363 | 3514792 | 3513718 | 3518530 | 3519691 | 3517577 | 3517560 | 3509694 | 3514405 | 3513583 | 3493571 | 3506588 | 3517657 | 3505506 | 3515579 | 3520464 | 3514300 | 3517475 | 3514183 | 3515794 | 3515696 | 3514615 | 3514596 | 3514382 | 3516703 | 3516232 | 55 | 3515210.6866 | 10070.5802 | 0.0028649
56 | 3077685 | 3078068 | 3076899 | 3069817 | 3027335 | 3074261 | 3078854 | 3078322 | 3044018 | 3073000 | 3075384 | 3076271 | 3075690 | 3075772 | 3069769 | 3074659 | 3047512 | 3065701 | 3073713 | 3075208 | 3075208 | 3076104 | 3080470 | 3062233 | 3050964 | 3077175 | 3075661 | 3081800 | 3074667 | 3075862 | 3080848 | 3080910 | 3079587 | 3077913 | 3063370 | 3072240 | 3077279 | 3075166 | 3067968 | 3072885 | 3072653 | 3076011 | 3072130 | 3073471 | 3069212 | 3070450 | 3067767 | 3069422 | 3069952 | 3068451 | 3066344 | 3050559 | 3068445 | 3070757 | 3062970 | 3067768 | 3074394 | 3072270 | 3074208 | 3069145 | 3072632 | 3069230 | 3069276 | 3071020 | 3069573 | 3070135 | 3074020 | 56 | 3070873.3284 | 9064.46 | 0.0029518
57 | 3356549 | 3354022 | 3357014 | 3346591 | 3301879 | 3348512 | 3353392 | 3354866 | 3322126 | 3350300 | 3355566 | 3353808 | 3349610 | 3353498 | 3349262 | 3353679 | 3324298 | 3345308 | 3352066 | 3350197 | 3350197 | 3351229 | 3350799 | 3340704 | 3328841 | 3352360 | 3360052 | 3355874 | 3352641 | 3355532 | 3355962 | 3357453 | 3354330 | 3353201 | 3338508 | 3350009 | 3353251 | 3348482 | 3342728 | 3351408 | 3353473 | 3349317 | 3351227 | 3349184 | 3348405 | 3349656 | 3351937 | 3347746 | 3345100 | 3347413 | 3345790 | 3329439 | 3341359 | 3347168 | 3342082 | 3345915 | 3351770 | 3346506 | 3351059 | 3350797 | 3348530 | 3346527 | 3348772 | 3347717 | 3344754 | 3347348 | 3348120 | 57 | 3347958.4328 | 9060.3577 | 0.0027062
58 | 3408537 | 3410250 | 3407391 | 3401566 | 3349745 | 3401770 | 3408229 | 3406606 | 3375410 | 3402169 | 3406010 | 3409324 | 3405484 | 3405881 | 3405995 | 3404670 | 3374847 | 3400371 | 3408204 | 3407138 | 3407138 | 3403325 | 3406839 | 3394507 | 3378335 | 3408050 | 3406389 | 3406772 | 3402735 | 3407383 | 3406839 | 3411040 | 3409901 | 3405451 | 3387469 | 3403570 | 3405964 | 3406065 | 3398947 | 3404971 | 3400597 | 3402257 | 3398077 | 3401370 | 3401303 | 3406043 | 3399492 | 3400389 | 3398077 | 3399357 | 3396860 | 3379116 | 3397394 | 3402579 | 3391746 | 3399503 | 3403964 | 3401562 | 3399284 | 3400285 | 3406961 | 3398125 | 3399963 | 3397279 | 3398606 | 3400440 | 3404067 | 58 | 3400686.3134 | 9860.1721 | 0.0028995
59 | 3494561 | 3490093 | 3487052 | 3489327 | 3441865 | 3486100 | 3500587 | 3496715 | 3456730 | 3480687 | 3492953 | 3499155 | 3492580 | 3489141 | 3486067 | 3490638 | 3459270 | 3476878 | 3486732 | 3489302 | 3489302 | 3483940 | 3485421 | 3472799 | 3465712 | 3486615 | 3488105 | 3491677 | 3483368 | 3489942 | 3488532 | 3488918 | 3492484 | 3486315 | 3471386 | 3484656 | 3488488 | 3485715 | 3477631 | 3488234 | 3487496 | 3479420 | 3480155 | 3482138 | 3479656 | 3481252 | 3479026 | 3477631 | 3475456 | 3478005 | 3474354 | 3460052 | 3476854 | 3479321 | 3473861 | 3478806 | 3480568 | 3480461 | 3479903 | 3481461 | 3480294 | 3478831 | 3481340 | 3481785 | 3480398 | 3482456 | 3482675 | 59 | 3482228.7761 | 9707.2707 | 0.0027877
60 | 3724478 | 3749606 | 3697980 | 3691413 | 3643814 | 3696316 | 3685131 | 3679844 | 3646859 | 3667591 | 3687499 | 3673198 | 3689802 | 3679372 | 3676037 | 3696521 | 3651756 | 3666897 | 3666119 | 3672150 | 3672150 | 3679955 | 3685775 | 3663988 | 3671815 | 3703756 | 3706265 | 3747926 | 3713194 | 3682005 | 3681890 | 3660040 | 3650210 | 3675785 | 3651625 | 3712042 | 3723187 | 3723143 | 3698463 | 3666832 | 3692076 | 3691886 | 3684185 | 3663795 | 3642532 | 3745214 | 3734237 | 3722561 | 3708341 | 3678359 | 3671563 | 3646221 | 3663819 | 3683844 | 3674954 | 3696121 | 3720516 | 3694358 | 3738782 | 3740986 | 3759533 | 3752845 | 3740227 | 3738120 | 3746001 | 3751042 | 3738443 | 60 | 3694522.2388 | 31509.2402 | 0.0085286
61 | 3456177 | 3450263 | 3455924 | 3449935 | 3397554 | 3446491 | 3455815 | 3452242 | 3422037 | 3450391 | 3455826 | 3452848 | 3449379 | 3448649 | 3449254 | 3453062 | 3420241 | 3444908 | 3452045 | 3453388 | 3453388 | 3451705 | 3450258 | 3440122 | 3425050 | 3451328 | 3455495 | 3454537 | 3451247 | 3451441 | 3453768 | 3455863 | 3452620 | 3450809 | 3433781 | 3446522 | 3455972 | 3448573 | 3445694 | 3452536 | 3449784 | 3451794 | 3446749 | 3452694 | 3451045 | 3451490 | 3450780 | 3445947 | 3449039 | 3446919 | 3446249 | 3426696 | 3441700 | 3448956 | 3440945 | 3447316 | 3449086 | 3447736 | 3446257 | 3446314 | 3453556 | 3447627 | 3449523 | 3446577 | 3445432 | 3445843 | 3450765 | 61 | 3447372.4925 | 9643.8692 | 0.0027975
62 | 3100219 | 3093763 | 3098088 | 3088499 | 3048307 | 3093809 | 3090093 | 3091419 | 3068096 | 3093133 | 3093805 | 3097389 | 3095122 | 3097400 | 3092111 | 3094105 | 3062976 | 3084319 | 3092075 | 3093974 | 3093974 | 3092475 | 3096079 | 3109597 | 3144721 | 3157191 | 3167874 | 3221734 | 3215534 | 3205101 | 3199083 | 3200818 | 3228273 | 3205699 | 3176398 | 3187980 | 3189766 | 3161564 | 3155977 | 3158063 | 3147455 | 3166319 | 3162531 | 3162139 | 3142123 | 3142109 | 3134031 | 3128022 | 3126853 | 3146212 | 3136303 | 3112583 | 3124969 | 3120715 | 3123614 | 3129239 | 3128826 | 3138847 | 3141364 | 3131642 | 3132266 | 3136649 | 3134494 | 3131755 | 3125099 | 3124016 | 3126190 | 62 | 3131238.3284 | 40092.627 | 0.012804
63 | 3752005 | 3750938 | 3749641 | 3747931 | 3686113 | 3748924 | 3751394 | 3754270 | 3713997 | 3747207 | 3754741 | 3751512 | 3746003 | 3751230 | 3747168 | 3748417 | 3714882 | 3741474 | 3752581 | 3750294 | 3750294 | 3752861 | 3748246 | 3735744 | 3722490 | 3752061 | 3751345 | 3750903 | 3751266 | 3749422 | 3748189 | 3754806 | 3748777 | 3750839 | 3727007 | 3742481 | 3752094 | 3742728 | 3739098 | 3747097 | 3745656 | 3746166 | 3747564 | 3748601 | 3741000 | 3745322 | 3744609 | 3740878 | 3743759 | 3742763 | 3742541 | 3722846 | 3737690 | 3743441 | 3737780 | 3741698 | 3749067 | 3746452 | 3743941 | 3743864 | 3745470 | 3743393 | 3745577 | 3744697 | 3741616 | 3742950 | 3745411 | 63 | 3743928.6866 | 10916.9275 | 0.0029159
64 | 2641286 | 2638809 | 2637219 | 2632479 | 2600573 | 2637315 | 2639149 | 2639002 | 2613776 | 2636193 | 2644285 | 2640849 | 2639867 | 2642025 | 2639637 | 2636355 | 2613757 | 2631126 | 2637529 | 2639814 | 2639814 | 2636124 | 2635924 | 2623775 | 2617527 | 2634460 | 2636393 | 2634898 | 2635215 | 2638620 | 2640339 | 2646663 | 2645022 | 2638143 | 2616711 | 2626413 | 2627800 | 2626256 | 2620630 | 2620160 | 2621388 | 2622747 | 2619296 | 2618940 | 2617861 | 2620542 | 2622704 | 2618309 | 2621238 | 2622810 | 2621469 | 2611202 | 2621228 | 2621764 | 2620071 | 2623550 | 2626963 | 2632747 | 2630880 | 2628949 | 2629163 | 2630417 | 2631795 | 2630285 | 2628804 | 2633713 | 2634016 | 64 | 2629623.6269 | 9508.5883 | 0.003616
65 | 2806462 | 2884275 | 2851193 | 2790400 | 2769248 | 2812313 | 2823101 | 2824308 | 2788690 | 2798920 | 2797434 | 2740004 | 2666566 | 2599792 | 2617244 | 2624247 | 2584303 | 2642794 | 2763716 | 2824164 | 2824164 | 2754584 | 2683233 | 2660126 | 2656953 | 2684474 | 2697657 | 2674245 | 2663850 | 2697883 | 2709176 | 2654217 | 2690255 | 2700768 | 2703639 | 2713000 | 2775926 | 2736085 | 2775704 | 2671978 | 2646495 | 2680190 | 2687645 | 2692500 | 2762324 | 2755314 | 2717796 | 2717484 | 2668963 | 2770533 | 2771961 | 2751133 | 2757641 | 2765744 | 2747184 | 2711162 | 2622543 | 2596108 | 2592139 | 2598363 | 2541943 | 2551127 | 2547952 | 2557244 | 2556010 | 2560623 | 2788153 | 65 | 2705274.1045 | 83180.3822 | 0.030747
66 | 3639292 | 3638754 | 3637952 | 3631697 | 3582988 | 3635619 | 3638757 | 3635857 | 3597792 | 3630998 | 3639863 | 3642304 | 3634366 | 3635792 | 3635560 | 3631124 | 3602894 | 3627960 | 3636179 | 3638969 | 3638969 | 3633396 | 3635339 | 3620972 | 3607036 | 3631190 | 3633956 | 3640559 | 3632979 | 3637711 | 3637213 | 3637388 | 3638052 | 3639097 | 3615752 | 3634026 | 3639662 | 3632878 | 3628188 | 3634665 | 3630428 | 3635371 | 3634315 | 3629199 | 3628280 | 3635221 | 3633382 | 3629909 | 3629379 | 3628425 | 3629959 | 3608828 | 3626181 | 3629468 | 3622344 | 3628703 | 3630792 | 3631971 | 3630782 | 3627906 | 3632518 | 3628230 | 3633002 | 3631266 | 3628734 | 3628573 | 3631740 | 66 | 3630427.6269 | 10205.1169 | 0.002811
67 | 3782444 | 3777875 | 3780927 | 3769056 | 3717917 | 3772213 | 3775255 | 3776912 | 3752697 | 3779853 | 3790178 | 3784128 | 3778481 | 3777741 | 3784622 | 3782450 | 3754204 | 3779814 | 3787612 | 3781825 | 3781825 | 3783361 | 3781071 | 3771193 | 3749695 | 3774443 | 3780404 | 3780726 | 3781968 | 3779715 | 3784654 | 3783729 | 3777418 | 3776518 | 3757535 | 3773890 | 3777456 | 3774882 | 3768126 | 3775481 | 3777483 | 3772804 | 3768304 | 3770591 | 3772366 | 3774884 | 3772315 | 3771472 | 3769049 | 3772459 | 3771834 | 3750247 | 3767828 | 3772527 | 3760940 | 3770243 | 3775278 | 3777239 | 3773065 | 3773440 | 3778410 | 3775249 | 3775064 | 3772626 | 3775626 | 3777515 | 3781100 | 67 | 3773884.3582 | 10636.3538 | 0.0028184
68 | 3614363 | 3611240 | 3613447 | 3606689 | 3551909 | 3607596 | 3614996 | 3616357 | 3575263 | 3606928 | 3614258 | 3611606 | 3609236 | 3610451 | 3608654 | 3612105 | 3578614 | 3605001 | 3611727 | 3611620 | 3611620 | 3610104 | 3611885 | 3597006 | 3584311 | 3611219 | 3611125 | 3614539 | 3609064 | 3616399 | 3614541 | 3616031 | 3612568 | 3610307 | 3595635 | 3612226 | 3613380 | 3608458 | 3603644 | 3612338 | 3610285 | 3612304 | 3608480 | 3609648 | 3605277 | 3609380 | 3607399 | 3611096 | 3605671 | 3605204 | 3606916 | 3586490 | 3603694 | 3604409 | 3594570 | 3609353 | 3607063 | 3610895 | 3605605 | 3607961 | 3604789 | 3602164 | 3610823 | 3607371 | 3606681 | 3604318 | 3606417 | 68 | 3606458.5522 | 10424.0708 | 0.0028904
69 | 3658088 | 3658438 | 3656016 | 3650104 | 3598238 | 3651688 | 3655786 | 3655148 | 3620503 | 3652129 | 3656700 | 3657682 | 3656872 | 3656671 | 3656417 | 3653265 | 3626745 | 3647580 | 3656912 | 3656642 | 3656642 | 3653794 | 3652020 | 3644388 | 3626744 | 3656082 | 3653683 | 3657355 | 3655750 | 3653562 | 3658331 | 3659585 | 3656083 | 3654484 | 3639806 | 3651465 | 3655347 | 3653693 | 3645900 | 3651184 | 3650256 | 3653070 | 3653206 | 3651343 | 3647653 | 3652091 | 3648859 | 3648051 | 3651391 | 3649720 | 3645460 | 3620134 | 3644269 | 3648895 | 3642270 | 3644194 | 3649754 | 3654367 | 3648262 | 3648200 | 3649727 | 3649689 | 3651786 | 3646716 | 3648944 | 3649657 | 3653792 | 69 | 3649541.4627 | 10196.2358 | 0.0027938
70 | 989212 | 986715 | 983120 | 979496 | 985175 | 992957 | 993224 | 988413 | 977343 | 984579 | 985889 | 988640 | 1010842 | 1060576 | 1057077 | 1045456 | 1058591 | 1057508 | 1045479 | 1040803 | 1040803 | 1042073 | 1030504 | 1022313 | 1009212 | 1016823 | 1012749 | 1011582 | 1006406 | 1012079 | 1018693 | 1018783 | 1010019 | 999834 | 991193 | 998944 | 986151 | 971811 | 974268 | 983736 | 981147 | 980232 | 978864 | 975470 | 979743 | 982038 | 982433 | 978232 | 967816 | 967214 | 965862 | 963194 | 967888 | 967792 | 967268 | 975340 | 982240 | 977691 | 981437 | 981640 | 978915 | 978459 | 986832 | 975256 | 964395 | 965043 | 969265 | 70 | 995384.7313 | 26627.8569 | 0.026751
71 | 3939600 | 3937352 | 3937323 | 3932023 | 3876247 | 3931932 | 3937872 | 3936338 | 3895822 | 3928686 | 3936989 | 3938113 | 3935091 | 3935744 | 3929807 | 3933556 | 3901194 | 3925129 | 3935499 | 3935133 | 3935133 | 3933415 | 3933253 | 3918305 | 3904458 | 3936031 | 3935421 | 3937520 | 3934255 | 3935888 | 3934815 | 3938640 | 3935888 | 3933195 | 3916976 | 3933578 | 3937552 | 3935518 | 3924751 | 3929256 | 3931314 | 3932857 | 3927106 | 3929577 | 3928634 | 3930607 | 3929202 | 3929003 | 3924397 | 3930408 | 3925354 | 3905419 | 3922264 | 3928648 | 3919986 | 3925614 | 3930359 | 3929914 | 3928837 | 3926953 | 3930398 | 3929118 | 3928655 | 3927187 | 3925125 | 3926797 | 3928171 | 71 | 3928585.1045 | 10723.47 | 0.0027296
72 | 3714500 | 3713337 | 3715677 | 3706405 | 3654701 | 3708085 | 3718007 | 3715417 | 3681772 | 3709712 | 3714395 | 3718207 | 3714441 | 3711848 | 3712938 | 3713016 | 3682014 | 3705693 | 3712364 | 3713107 | 3713107 | 3710997 | 3714342 | 3693924 | 3686021 | 3714532 | 3712717 | 3715072 | 3712592 | 3711590 | 3711978 | 3714321 | 3719374 | 3713122 | 3696234 | 3709706 | 3713434 | 3713167 | 3705549 | 3711462 | 3711404 | 3712640 | 3709040 | 3713982 | 3706651 | 3708304 | 3706572 | 3704249 | 3708000 | 3706371 | 3704265 | 3684712 | 3697271 | 3708134 | 3697907 | 3703514 | 3709701 | 3707650 | 3703526 | 3703450 | 3709458 | 3706791 | 3704817 | 3707458 | 3708523 | 3709190 | 3708815 | 72 | 3707332.4179 | 10252.2554 | 0.0027654
73 | 2608797 | 2610280 | 2606726 | 2604965 | 2570092 | 2607582 | 2608241 | 2611758 | 2582511 | 2605319 | 2610524 | 2606189 | 2607244 | 2611312 | 2603382 | 2604001 | 2582561 | 2602516 | 2607811 | 2605046 | 2605046 | 2605688 | 2606216 | 2593941 | 2587290 | 2607644 | 2605430 | 2604853 | 2602938 | 2603948 | 2605946 | 2604737 | 2606363 | 2605891 | 2593649 | 2605781 | 2612605 | 2609125 | 2602856 | 2604420 | 2604471 | 2606668 | 2602594 | 2605797 | 2602382 | 2605444 | 2601436 | 2600121 | 2601087 | 2598673 | 2595623 | 2583672 | 2591422 | 2598338 | 2593498 | 2598132 | 2602221 | 2599457 | 2600602 | 2600300 | 2597788 | 2594979 | 2597819 | 2600010 | 2599174 | 2600779 | 2603950 | 73 | 2601756.1343 | 7540.8232 | 0.0028984
74 | 3790737 | 3784781 | 3783271 | 3778870 | 3728291 | 3782920 | 3790040 | 3785302 | 3748442 | 3781119 | 3787936 | 3784890 | 3785899 | 3781312 | 3780813 | 3781577 | 3753569 | 3773540 | 3780792 | 3783210 | 3783210 | 3781857 | 3778448 | 3770862 | 3756282 | 3781842 | 3785685 | 3785363 | 3780802 | 3780748 | 3780520 | 3779504 | 3782090 | 3778300 | 3762057 | 3778695 | 3782608 | 3783666 | 3773981 | 3780562 | 3776914 | 3778943 | 3781180 | 3780180 | 3775327 | 3780385 | 3778669 | 3773622 | 3774563 | 3774797 | 3771097 | 3746630 | 3771196 | 3774700 | 3767448 | 3774510 | 3780439 | 3773146 | 3772437 | 3775005 | 3772510 | 3772724 | 3776046 | 3770882 | 3774069 | 3776858 | 3777175 | 74 | 3776355.8955 | 10273.9616 | 0.0027206
75 | 3909373 | 3908206 | 3904437 | 3893527 | 3842813 | 3901932 | 3903116 | 3909313 | 3864180 | 3899197 | 3905724 | 3904655 | 3905541 | 3901463 | 3900032 | 3902296 | 3869195 | 3892600 | 3905163 | 3903788 | 3903788 | 3902871 | 3901469 | 3887290 | 3874964 | 3902212 | 3903113 | 3908353 | 3902435 | 3902140 | 3901910 | 3903650 | 3905662 | 3904196 | 3883586 | 3901767 | 3905091 | 3902863 | 3895291 | 3899292 | 3901847 | 3899899 | 3897614 | 3900456 | 3896516 | 3898287 | 3893131 | 3895122 | 3894301 | 3895631 | 3894352 | 3872289 | 3891651 | 3897047 | 3889550 | 3889509 | 3896663 | 3894334 | 3896685 | 3895416 | 3894701 | 3895159 | 3893824 | 3897515 | 3899206 | 3899401 | 3896612 | 75 | 3896794.209 | 10962.3634 | 0.0028132
76 | 3930980 | 3932693 | 3933221 | 3923307 | 3868039 | 3924274 | 3931596 | 3928477 | 3894688 | 3925821 | 3934592 | 3937191 | 3928514 | 3935147 | 3925053 | 3929207 | 3895896 | 3924366 | 3930995 | 3930017 | 3930017 | 3930694 | 3923485 | 3910856 | 3897388 | 3931291 | 3928914 | 3933587 | 3928128 | 3926183 | 3930635 | 3932278 | 3932748 | 3930476 | 3911124 | 3929598 | 3930172 | 3931852 | 3921290 | 3924066 | 3926066 | 3929652 | 3925527 | 3928927 | 3923566 | 3927817 | 3923941 | 3923328 | 3921709 | 3921969 | 3921574 | 3896481 | 3917602 | 3923884 | 3913738 | 3922109 | 3928440 | 3923595 | 3920437 | 3921126 | 3921755 | 3920941 | 3924512 | 3919427 | 3920151 | 3923930 | 3927249 | 76 | 3923556.8507 | 11192.6504 | 0.0028527
77 | 3794647 | 3800229 | 3798964 | 3794751 | 3736307 | 3792143 | 3800649 | 3799556 | 3761600 | 3794972 | 3798841 | 3799253 | 3795759 | 3799954 | 3793981 | 3791966 | 3764546 | 3790477 | 3798670 | 3797779 | 3797779 | 3794173 | 3792978 | 3779167 | 3765731 | 3796895 | 3799170 | 3800911 | 3796473 | 3793628 | 3795473 | 3800445 | 3795956 | 3792099 | 3776867 | 3787994 | 3797113 | 3792395 | 3789095 | 3792297 | 3795054 | 3790587 | 3793936 | 3793133 | 3793347 | 3796738 | 3787766 | 3792572 | 3787187 | 3789467 | 3788142 | 3765535 | 3781562 | 3789218 | 3781626 | 3785685 | 3790834 | 3788567 | 3791242 | 3789264 | 3791807 | 3788427 | 3788237 | 3787709 | 3783969 | 3787842 | 3791443 | 77 | 3790038.4925 | 10734.9757 | 0.0028324
78 | 4041549 | 4043355 | 4038710 | 4028821 | 3974475 | 4036762 | 4037356 | 4039175 | 4001507 | 4034318 | 4044772 | 4040570 | 4037505 | 4037002 | 4029593 | 4032566 | 4006982 | 4031316 | 4040455 | 4038501 | 4038501 | 4037447 | 4035915 | 4019894 | 4003862 | 4036429 | 4035657 | 4042028 | 4036624 | 4038274 | 4039200 | 4041197 | 4039225 | 4037470 | 4018624 | 4034690 | 4040390 | 4034031 | 4027462 | 4033227 | 4035331 | 4037633 | 4031723 | 4035271 | 4033870 | 4029596 | 4030998 | 4032347 | 4031768 | 4027630 | 4026010 | 4004307 | 4023660 | 4031160 | 4017993 | 4026173 | 4035707 | 4033720 | 4027954 | 4026458 | 4032937 | 4028224 | 4030206 | 4030437 | 4029165 | 4029547 | 4034299 | 78 | 4031038.2239 | 11336.1181 | 0.0028122
79 | 3945015 | 3943964 | 3941892 | 3938009 | 3878464 | 3939114 | 3941982 | 3944457 | 3903055 | 3936018 | 3942305 | 3943230 | 3942475 | 3939911 | 3938511 | 3938801 | 3909556 | 3936876 | 3941185 | 3942874 | 3942874 | 3940984 | 3939850 | 3927507 | 3910099 | 3940032 | 3946571 | 3941595 | 3939580 | 3940403 | 3938653 | 3943156 | 3946470 | 3938023 | 3924668 | 3940137 | 3944836 | 3941083 | 3934700 | 3933190 | 3938192 | 3936019 | 3937302 | 3936733 | 3931828 | 3935808 | 3934714 | 3929610 | 3933940 | 3936434 | 3930596 | 3906615 | 3929188 | 3934527 | 3926331 | 3935374 | 3937677 | 3936070 | 3934883 | 3934156 | 3937031 | 3933422 | 3932482 | 3933479 | 3930431 | 3933743 | 3935034 | 79 | 3934831.7015 | 11044.1695 | 0.0028068
80 | 4123291 | 4122937 | 4122205 | 4115471 | 4058976 | 4116579 | 4125543 | 4123283 | 4084350 | 4114235 | 4126228 | 4122550 | 4121025 | 4122851 | 4120754 | 4117813 | 4089624 | 4113599 | 4123941 | 4120524 | 4120524 | 4118602 | 4118648 | 4103997 | 4088773 | 4120263 | 4120298 | 4124468 | 4115698 | 4118058 | 4122425 | 4122432 | 4120498 | 4119695 | 4101868 | 4113435 | 4121691 | 4118901 | 4110892 | 4115610 | 4115113 | 4115598 | 4116138 | 4117902 | 4110455 | 4115383 | 4112049 | 4112679 | 4110963 | 4112682 | 4109772 | 4086443 | 4108509 | 4112644 | 4102681 | 4106276 | 4117376 | 4115103 | 4109211 | 4110214 | 4114656 | 4114388 | 4114404 | 4109187 | 4111021 | 4111694 | 4116075 | 80 | 4113599.5672 | 11020.4847 | 0.002679
81 | 4559480 | 4562602 | 4567631 | 4557247 | 4494482 | 4561414 | 4564867 | 4565940 | 4519314 | 4551126 | 4565203 | 4567661 | 4562345 | 4562908 | 4562477 | 4558961 | 4524801 | 4557365 | 4563518 | 4561151 | 4561151 | 4564049 | 4559308 | 4543250 | 4526643 | 4554436 | 4555690 | 4558820 | 4557894 | 4554931 | 4561284 | 4558487 | 4559649 | 4559488 | 4536138 | 4556488 | 4563240 | 4557561 | 4549701 | 4555137 | 4556230 | 4558834 | 4553421 | 4555902 | 4549478 | 4559158 | 4555700 | 4553167 | 4554784 | 4555415 | 4548977 | 4525396 | 4549318 | 4551942 | 4543264 | 4553617 | 4559714 | 4552534 | 4551979 | 4558180 | 4552690 | 4553711 | 4552364 | 4558271 | 4554123 | 4552875 | 4562083 | 81 | 4553984.5522 | 12102.8307 | 0.0026576
82 | 3937503 | 3936607 | 3940313 | 3929808 | 3877612 | 3934982 | 3938924 | 3940223 | 3900446 | 3932526 | 3939185 | 3935848 | 3935853 | 3937215 | 3935096 | 3935011 | 3907002 | 3929497 | 3935852 | 3936804 | 3936804 | 3936811 | 3934730 | 3924086 | 3908161 | 3934932 | 3938364 | 3936540 | 3934624 | 3934926 | 3937156 | 3937462 | 3937300 | 3935298 | 3922832 | 3931796 | 3943916 | 3935490 | 3930243 | 3933192 | 3934068 | 3937707 | 3932687 | 3936621 | 3932211 | 3933595 | 3931848 | 3929049 | 3926620 | 3927833 | 3931546 | 3904349 | 3921741 | 3930583 | 3919953 | 3926153 | 3931671 | 3932699 | 3932189 | 3931868 | 3933380 | 3927772 | 3931220 | 3929545 | 3928181 | 3933435 | 3932850 | 82 | 3930781.2537 | 10381.2063 | 0.002641
83 | 4497965 | 4495490 | 4502426 | 4492150 | 4424817 | 4488102 | 4497155 | 4497674 | 4457516 | 4488194 | 4496595 | 4493846 | 4494413 | 4495515 | 4493153 | 4490128 | 4455936 | 4486895 | 4503974 | 4496308 | 4496308 | 4497506 | 4493206 | 4475892 | 4465412 | 4493325 | 4495862 | 4499696 | 4490248 | 4495188 | 4496893 | 4499425 | 4495011 | 4492865 | 4474553 | 4493994 | 4491464 | 4492542 | 4486005 | 4492882 | 4490116 | 4490610 | 4492797 | 4495112 | 4490388 | 4489018 | 4488671 | 4485086 | 4484785 | 4489818 | 4485818 | 4461935 | 4479373 | 4489921 | 4481418 | 4485533 | 4489169 | 4491346 | 4487664 | 4485638 | 4486837 | 4486127 | 4486498 | 4484462 | 4486486 | 4487967 | 4492026 | 83 | 4488226.0896 | 12054.1425 | 0.0026857
84 | 4513478 | 4510614 | 4508545 | 4498581 | 4441832 | 4504842 | 4511199 | 4509156 | 4470703 | 4500325 | 4510249 | 4512751 | 4505568 | 4508295 | 4503162 | 4503001 | 4470518 | 4496578 | 4508633 | 4511476 | 4511476 | 4506899 | 4505209 | 4492929 | 4471503 | 4506618 | 4508029 | 4510875 | 4503997 | 4507622 | 4507894 | 4509198 | 4506693 | 4508127 | 4483636 | 4501796 | 4503962 | 4507158 | 4499124 | 4504766 | 4502479 | 4503981 | 4501560 | 4501687 | 4501296 | 4503984 | 4500144 | 4498237 | 4494087 | 4497189 | 4493217 | 4473436 | 4488825 | 4496180 | 4488975 | 4496858 | 4503366 | 4501462 | 4498212 | 4498013 | 4499479 | 4495042 | 4496745 | 4499098 | 4497587 | 4497764 | 4502435 | 84 | 4499826.194 | 11901.8424 | 0.002645
85 | 4128180 | 4126778 | 4128031 | 4124014 | 4064965 | 4123886 | 4128899 | 4126738 | 4087577 | 4119528 | 4131268 | 4131264 | 4125133 | 4129407 | 4124275 | 4124414 | 4095969 | 4119274 | 4126088 | 4123760 | 4123760 | 4125050 | 4123247 | 4112047 | 4096239 | 4123031 | 4129423 | 4129353 | 4121604 | 4127321 | 4127612 | 4128522 | 4127552 | 4127193 | 4106400 | 4121087 | 4124762 | 4123400 | 4116329 | 4124815 | 4123833 | 4125168 | 4119606 | 4121310 | 4122340 | 4122979 | 4119914 | 4115818 | 4119941 | 4116094 | 4115118 | 4087312 | 4110649 | 4118417 | 4106116 | 4115036 | 4119838 | 4117430 | 4117671 | 4120219 | 4118991 | 4116911 | 4115927 | 4120453 | 4118063 | 4117057 | 4117890 | 85 | 4119228.2985 | 11178.1116 | 0.0027136
86 | 5061704 | 5061850 | 5059606 | 5047298 | 4977549 | 5051992 | 5059250 | 5053578 | 5009369 | 5049363 | 5059150 | 5057710 | 5056375 | 5054788 | 5053921 | 5052954 | 5007718 | 5046339 | 5052718 | 5055806 | 5055806 | 5054228 | 5052989 | 5028845 | 5022703 | 5051496 | 5059278 | 5056559 | 5056096 | 5055487 | 5056699 | 5058441 | 5056133 | 5054635 | 5031905 | 5048167 | 5054788 | 5052993 | 5045471 | 5049866 | 5048340 | 5052427 | 5050126 | 5054906 | 5044710 | 5050431 | 5045803 | 5046981 | 5044085 | 5042411 | 5038372 | 5012917 | 5038722 | 5045800 | 5035074 | 5041647 | 5049035 | 5047192 | 5048671 | 5046638 | 5044979 | 5043276 | 5046896 | 5045535 | 5045352 | 5045487 | 5048539 | 86 | 5046805.597 | 13968.6873 | 0.0027678
87 | 4744891 | 4738599 | 4740082 | 4734545 | 4668364 | 4736171 | 4739415 | 4743201 | 4694943 | 4731773 | 4744624 | 4742495 | 4737663 | 4739726 | 4735780 | 4731715 | 4697054 | 4730723 | 4740751 | 4739635 | 4739635 | 4737629 | 4733114 | 4719972 | 4704058 | 4739340 | 4737750 | 4740920 | 4739174 | 4738374 | 4737573 | 4743952 | 4737603 | 4735051 | 4714236 | 4730248 | 4736879 | 4733113 | 4729467 | 4734804 | 4734751 | 4736602 | 4730861 | 4732823 | 4732066 | 4737391 | 4732394 | 4728781 | 4724929 | 4726913 | 4726654 | 4701582 | 4726519 | 4729533 | 4721085 | 4728094 | 4734524 | 4732375 | 4727262 | 4727086 | 4731947 | 4725175 | 4724953 | 4727243 | 4728848 | 4727504 | 4734026 | 87 | 4730551.6866 | 12707.7268 | 0.0026863
88 | 5332556 | 5326405 | 5325449 | 5317888 | 5251556 | 5319491 | 5332135 | 5329177 | 5277328 | 5320567 | 5332951 | 5328389 | 5322564 | 5328811 | 5321323 | 5321903 | 5282226 | 5316620 | 5324448 | 5331348 | 5331348 | 5325677 | 5317597 | 5306291 | 5286829 | 5321892 | 5325536 | 5328256 | 5323996 | 5324121 | 5329115 | 5331560 | 5327709 | 5328195 | 5305079 | 5319545 | 5330612 | 5324565 | 5315995 | 5320141 | 5324568 | 5320218 | 5321178 | 5318739 | 5315140 | 5325076 | 5320187 | 5316552 | 5316947 | 5318619 | 5315420 | 5285182 | 5309462 | 5317131 | 5299881 | 5312285 | 5324651 | 5322112 | 5316563 | 5316161 | 5316165 | 5315346 | 5316197 | 5315866 | 5316305 | 5318177 | 5319364 | 88 | 5317920.6866 | 13903.9793 | 0.0026146
89 | 4901483 | 4902997 | 4902976 | 4893187 | 4828942 | 4894650 | 4897136 | 4900982 | 4856628 | 4896858 | 4905745 | 4906150 | 4896086 | 4899884 | 4895471 | 4894518 | 4857205 | 4889843 | 4901171 | 4898806 | 4898806 | 4894597 | 4894374 | 4876258 | 4866292 | 4899262 | 4901433 | 4901463 | 4896314 | 4898142 | 4896723 | 4894921 | 4900521 | 4899991 | 4869162 | 4889515 | 4898327 | 4891273 | 4886071 | 4890626 | 4895270 | 4895825 | 4896279 | 4891025 | 4888963 | 4895858 | 4893476 | 4891941 | 4887111 | 4888631 | 4883390 | 4857763 | 4880908 | 4885003 | 4874794 | 4884570 | 4889027 | 4887959 | 4881366 | 4886245 | 4888430 | 4885242 | 4888000 | 4885780 | 4884192 | 4887663 | 4889910 | 89 | 4889692.6866 | 13070.0414 | 0.002673
90 | 5482671 | 5483283 | 5485483 | 5469888 | 5398348 | 5479662 | 5480774 | 5480012 | 5432313 | 5477444 | 5488026 | 5483321 | 5475600 | 5480151 | 5474787 | 5479140 | 5433592 | 5474699 | 5481333 | 5477997 | 5477997 | 5474808 | 5475669 | 5459128 | 5438809 | 5477588 | 5479517 | 5481576 | 5477182 | 5476040 | 5482507 | 5481848 | 5483152 | 5476771 | 5452148 | 5474196 | 5477259 | 5477709 | 5467606 | 5472032 | 5472880 | 5475561 | 5473627 | 5471930 | 5462728 | 5474941 | 5473504 | 5466258 | 5465662 | 5465544 | 5468642 | 5439022 | 5459132 | 5469216 | 5455451 | 5466562 | 5473170 | 5471108 | 5466783 | 5471696 | 5469144 | 5463433 | 5469934 | 5467202 | 5465118 | 5468086 | 5471119 | 90 | 5470142.0746 | 14423.984 | 0.0026369
91 | 3672611 | 3679323 | 3681086 | 3669516 | 3615954 | 3669541 | 3680197 | 3677385 | 3639339 | 3673786 | 3678891 | 3676917 | 3673157 | 3676736 | 3676614 | 3672581 | 3642710 | 3669268 | 3678658 | 3678205 | 3678205 | 3675546 | 3673722 | 3662989 | 3647064 | 3675230 | 3676250 | 3678783 | 3675201 | 3674002 | 3680801 | 3680021 | 3675756 | 3675762 | 3657884 | 3670095 | 3678198 | 3674490 | 3667072 | 3674149 | 3671891 | 3674128 | 3668320 | 3669010 | 3670607 | 3672233 | 3672606 | 3667698 | 3669192 | 3666625 | 3665601 | 3652376 | 3661736 | 3669989 | 3662868 | 3667998 | 3671455 | 3673924 | 3670943 | 3669608 | 3672573 | 3671372 | 3673715 | 3669774 | 3671980 | 3669173 | 3674129 | 91 | 3670227.1493 | 10477.0455 | 0.0028546
92 | 3429856 | 3425384 | 3427089 | 3423627 | 3370172 | 3418373 | 3429974 | 3423472 | 3391375 | 3421513 | 3425626 | 3422080 | 3422948 | 3423477 | 3423535 | 3423837 | 3392758 | 3417013 | 3427215 | 3424976 | 3424976 | 3425230 | 3423088 | 3409838 | 3400242 | 3427307 | 3426319 | 3428360 | 3426578 | 3427348 | 3429946 | 3428391 | 3430775 | 3423448 | 3409723 | 3425927 | 3428699 | 3422895 | 3417857 | 3422778 | 3423128 | 3422895 | 3420847 | 3421242 | 3418143 | 3420662 | 3420594 | 3417262 | 3421082 | 3418157 | 3421751 | 3397839 | 3414058 | 3422818 | 3411792 | 3423634 | 3424169 | 3420999 | 3416946 | 3420984 | 3419883 | 3419331 | 3418209 | 3420722 | 3417222 | 3418497 | 3420890 | 92 | 3419966.8806 | 9901.4252 | 0.0028952
93 | 2947571 | 2947198 | 2944618 | 2938809 | 2900093 | 2940896 | 2945154 | 2947306 | 2917348 | 2939570 | 2943988 | 2945905 | 2943654 | 2940719 | 2945581 | 2943765 | 2920796 | 2943949 | 2950193 | 2949734 | 2949734 | 2943232 | 2943131 | 2934254 | 2925826 | 2947734 | 2942099 | 2944655 | 2944614 | 2946406 | 2948421 | 2952361 | 2952758 | 2949228 | 2935878 | 2942302 | 2943378 | 2937499 | 2934657 | 2940487 | 2936638 | 2938976 | 2937157 | 2937141 | 2930271 | 2936080 | 2936751 | 2934200 | 2933191 | 2934522 | 2932777 | 2913764 | 2926500 | 2933836 | 2930198 | 2931228 | 2938697 | 2937271 | 2940200 | 2938402 | 2938779 | 2940144 | 2937968 | 2939510 | 2939772 | 2939324 | 2939472 | 93 | 2938929.403 | 8963.9177 | 0.0030501
94 | 3503165 | 3504467 | 3498640 | 3497080 | 3444879 | 3496633 | 3502413 | 3501287 | 3465781 | 3503349 | 3501437 | 3503359 | 3496061 | 3502208 | 3499258 | 3507247 | 3473999 | 3495876 | 3502500 | 3505239 | 3505239 | 3504115 | 3501214 | 3484222 | 3477311 | 3500930 | 3501071 | 3504539 | 3501394 | 3503767 | 3504517 | 3509106 | 3506198 | 3504272 | 3485933 | 3500886 | 3503283 | 3502137 | 3491022 | 3500365 | 3499499 | 3501064 | 3493753 | 3497903 | 3498239 | 3506326 | 3499979 | 3501680 | 3500514 | 3498921 | 3498187 | 3478343 | 3498266 | 3501694 | 3492570 | 3498373 | 3503762 | 3502639 | 3501734 | 3501600 | 3504194 | 3503932 | 3500728 | 3502010 | 3501711 | 3502583 | 3503830 | 94 | 3498364.6716 | 10173.0502 | 0.0029079
95 | 3106432 | 3106668 | 3102933 | 3100556 | 3054465 | 3100833 | 3108532 | 3111321 | 3074373 | 3103777 | 3108403 | 3109042 | 3102496 | 3105251 | 3100660 | 3103130 | 3077732 | 3099540 | 3103527 | 3104953 | 3104953 | 3102060 | 3098550 | 3088757 | 3080122 | 3099031 | 3100071 | 3099502 | 3105015 | 3102986 | 3103346 | 3103996 | 3102756 | 3097671 | 3088161 | 3097822 | 3096555 | 3095539 | 3088881 | 3093363 | 3092907 | 3091640 | 3090640 | 3092584 | 3089629 | 3091819 | 3090555 | 3091697 | 3089238 | 3091071 | 3090609 | 3073350 | 3088083 | 3097942 | 3085212 | 3093999 | 3097165 | 3098068 | 3095465 | 3099440 | 3100786 | 3098719 | 3094386 | 3099907 | 3096323 | 3099292 | 3100880 | 95 | 3096345.7761 | 9399.4408 | 0.0030357
96 | 3913047 | 3914691 | 3908000 | 3902538 | 3842489 | 3906391 | 3909485 | 3907211 | 3869844 | 3899632 | 3904568 | 3900937 | 3901743 | 3903239 | 3894806 | 3897888 | 3876004 | 3890922 | 3903617 | 3904634 | 3904634 | 3903740 | 3903771 | 3886953 | 3880725 | 3901955 | 3902654 | 3904590 | 3903459 | 3904709 | 3905150 | 3910753 | 3910473 | 3907384 | 3886014 | 3897922 | 3911534 | 3904823 | 3893683 | 3902086 | 3904548 | 3904337 | 3906879 | 3901530 | 3900775 | 3903631 | 3898501 | 3897749 | 3901502 | 3898402 | 3896843 | 3873230 | 3892299 | 3892412 | 3884198 | 3890031 | 3907786 | 3901998 | 3903832 | 3905031 | 3906146 | 3902947 | 3905226 | 3903255 | 3901035 | 3905247 | 3905113 | 96 | 3899599.7164 | 11277.2214 | 0.0028919
97 | 933554 | 915103 | 904940 | 907365 | 899052 | 945977 | 916490 | 946727 | 968008 | 923656 | 902745 | 917093 | 926835 | 904477 | 869493 | 885102 | 880475 | 901488 | 927621 | 934841 | 934841 | 967410 | 982480 | 955276 | 936443 | 936133 | 911069 | 926451 | 922261 | 911985 | 913537 | 921980 | 922184 | 909837 | 938149 | 943048 | 936900 | 943637 | 942710 | 957304 | 965901 | 961552 | 940700 | 952129 | 939804 | 922424 | 889171 | 913696 | 909798 | 912450 | 886664 | 906495 | 931808 | 957736 | 937729 | 930855 | 935555 | 942245 | 942651 | 924839 | 925814 | 919049 | 907724 | 906839 | 903836 | 898123 | 907567 | 97 | 925340.7612 | 22576.5898 | 0.024398
98 | 3464273 | 3456785 | 3460010 | 3450941 | 3401896 | 3452363 | 3462328 | 3457003 | 3423486 | 3451811 | 3459200 | 3462931 | 3458624 | 3458428 | 3456010 | 3456579 | 3425338 | 3447103 | 3457196 | 3457401 | 3457401 | 3455642 | 3457631 | 3441080 | 3429230 | 3458070 | 3457276 | 3457733 | 3459947 | 3455290 | 3459631 | 3460992 | 3458904 | 3458898 | 3438312 | 3452362 | 3460924 | 3454006 | 3448839 | 3456145 | 3451967 | 3454505 | 3452473 | 3452866 | 3455786 | 3457588 | 3451847 | 3452994 | 3449750 | 3451062 | 3449401 | 3429340 | 3446024 | 3452916 | 3442234 | 3450456 | 3452241 | 3452754 | 3452301 | 3451556 | 3457897 | 3451669 | 3454797 | 3451489 | 3449744 | 3451425 | 3453545 | 98 | 3451920.0896 | 10192.4247 | 0.0029527
99 | 3547955 | 3545693 | 3544971 | 3537267 | 3485773 | 3533912 | 3543582 | 3542896 | 3508921 | 3536606 | 3544428 | 3545299 | 3540749 | 3541700 | 3536762 | 3538498 | 3509248 | 3534916 | 3540530 | 3542615 | 3542615 | 3544467 | 3537590 | 3528402 | 3515096 | 3539956 | 3541618 | 3542570 | 3540348 | 3537929 | 3539836 | 3542485 | 3543122 | 3541725 | 3521378 | 3537544 | 3543586 | 3541255 | 3533806 | 3538262 | 3538945 | 3540594 | 3536742 | 3537460 | 3535284 | 3543055 | 3537661 | 3535769 | 3536575 | 3536670 | 3531622 | 3514120 | 3530251 | 3536076 | 3526571 | 3535084 | 3540944 | 3540180 | 3534878 | 3533548 | 3535478 | 3533349 | 3541084 | 3536226 | 3537791 | 3534144 | 3536498 | 99 | 3536007.6119 | 9969.2058 | 0.0028193
100 | 2984282 | 2982312 | 2984402 | 2974146 | 2935308 | 2976518 | 2981752 | 2981763 | 2953577 | 2978483 | 2981479 | 2983329 | 2978763 | 2984205 | 2974932 | 2982917 | 2957360 | 2972038 | 2978971 | 2981236 | 2981236 | 2976642 | 2978292 | 2967504 | 2958775 | 2980143 | 2979023 | 2978520 | 2977803 | 2978386 | 2981604 | 2979987 | 2978824 | 2982252 | 2966312 | 2977848 | 2978336 | 2979107 | 2972116 | 2976820 | 2977180 | 2978156 | 2975043 | 2973367 | 2977019 | 2977374 | 2973071 | 2969944 | 2971997 | 2970533 | 2971676 | 2951070 | 2965696 | 2974204 | 2965028 | 2971415 | 2972385 | 2973710 | 2971276 | 2971205 | 2972374 | 2970426 | 2970723 | 2973622 | 2969834 | 2977148 | 2973329 | 100 | 2974270.2687 | 8351.6848 | 0.002808
101 | 3278832 | 3276194 | 3280271 | 3270676 | 3222061 | 3271363 | 3277607 | 3278420 | 3247462 | 3268830 | 3282300 | 3275344 | 3274958 | 3278144 | 3274096 | 3273211 | 3249455 | 3272624 | 3278773 | 3272117 | 3272117 | 3275105 | 3271705 | 3261136 | 3245854 | 3275683 | 3275744 | 3279991 | 3270833 | 3272835 | 3274172 | 3274542 | 3273985 | 3272775 | 3257271 | 3274986 | 3280862 | 3270513 | 3269793 | 3274430 | 3271990 | 3272721 | 3271418 | 3270653 | 3267106 | 3274276 | 3270022 | 3270254 | 3268472 | 3268908 | 3264491 | 3249729 | 3263932 | 3269092 | 3254972 | 3268852 | 3268055 | 3267673 | 3267584 | 3270372 | 3268604 | 3266051 | 3269102 | 3266179 | 3263280 | 3272174 | 3268771 | 101 | 3269429.8955 | 9504.014 | 0.0029069
102 | 2961542 | 2963437 | 2960634 | 2957512 | 2917680 | 2960979 | 2963425 | 2962384 | 2931022 | 2957404 | 2963839 | 2963720 | 2961005 | 2961727 | 2955353 | 2959343 | 2937262 | 2951380 | 2962170 | 2959491 | 2959491 | 2959101 | 2960988 | 2945654 | 2941216 | 2959036 | 2961569 | 2961222 | 2958133 | 2952827 | 2960696 | 2960194 | 2958057 | 2958782 | 2942044 | 2954308 | 2960777 | 2959376 | 2953864 | 2957647 | 2956381 | 2956098 | 2954372 | 2956156 | 2956936 | 2956335 | 2950105 | 2952995 | 2949547 | 2951043 | 2949202 | 2933361 | 2950173 | 2950466 | 2948887 | 2947230 | 2953715 | 2953980 | 2951636 | 2953547 | 2949959 | 2949019 | 2953019 | 2949393 | 2953438 | 2952729 | 2956455 | 102 | 2954215.9403 | 8217.28 | 0.0027815
103 | 3496027 | 3496276 | 3491548 | 3490494 | 3436883 | 3492870 | 3492102 | 3494726 | 3460869 | 3489748 | 3496188 | 3499416 | 3494954 | 3494796 | 3491593 | 3491543 | 3459704 | 3487364 | 3492479 | 3494118 | 3494118 | 3495036 | 3490096 | 3477392 | 3467504 | 3492564 | 3493433 | 3497370 | 3492267 | 3498158 | 3496982 | 3495460 | 3490638 | 3489948 | 3477492 | 3490425 | 3495690 | 3490641 | 3480756 | 3492757 | 3490108 | 3489470 | 3492083 | 3491269 | 3488647 | 3490308 | 3487020 | 3487515 | 3488178 | 3487705 | 3484506 | 3464938 | 3483088 | 3490636 | 3480177 | 3483979 | 3489420 | 3488863 | 3491133 | 3489087 | 3489865 | 3486076 | 3489516 | 3489997 | 3484567 | 3485883 | 3493004 | 103 | 3488051.6866 | 10085.8281 | 0.0028915
104 | 3543386 | 3540944 | 3539435 | 3533244 | 3484864 | 3540910 | 3543146 | 3540324 | 3505542 | 3538469 | 3540931 | 3538335 | 3540643 | 3542287 | 3535467 | 3539945 | 3508609 | 3533019 | 3543734 | 3543685 | 3543685 | 3540311 | 3539735 | 3522854 | 3513786 | 3537771 | 3538808 | 3539254 | 3535382 | 3542104 | 3541177 | 3540799 | 3539495 | 3537902 | 3522832 | 3536920 | 3539085 | 3538491 | 3533535 | 3538227 | 3540240 | 3539916 | 3534485 | 3537621 | 3533120 | 3539290 | 3539604 | 3533892 | 3533347 | 3532319 | 3534848 | 3512072 | 3530528 | 3535143 | 3526886 | 3532963 | 3533548 | 3535030 | 3532533 | 3531724 | 3534415 | 3528143 | 3536308 | 3532204 | 3531649 | 3537321 | 3537729 | 104 | 3534417.1642 | 9905.1326 | 0.0028025
105 | 3481792 | 3480171 | 3479957 | 3475130 | 3421878 | 3473728 | 3480985 | 3483245 | 3446103 | 3475864 | 3477334 | 3480459 | 3480321 | 3476502 | 3476884 | 3476359 | 3450104 | 3471291 | 3480684 | 3482921 | 3482921 | 3479626 | 3478678 | 3463997 | 3453591 | 3477094 | 3481250 | 3481839 | 3481234 | 3480358 | 3483868 | 3481647 | 3479168 | 3479452 | 3460440 | 3478287 | 3481549 | 3477887 | 3471189 | 3476799 | 3477995 | 3474582 | 3475322 | 3473645 | 3475274 | 3478266 | 3472625 | 3473855 | 3478589 | 3471925 | 3469109 | 3454089 | 3467858 | 3472217 | 3463001 | 3470275 | 3479175 | 3474323 | 3476269 | 3473022 | 3473093 | 3473947 | 3476626 | 3473616 | 3475447 | 3476094 | 3475525 | 105 | 3474005.2239 | 9972.6035 | 0.0028706
106 | 3490023 | 3487860 | 3487233 | 3486172 | 3435424 | 3485230 | 3491848 | 3491347 | 3456714 | 3487321 | 3486161 | 3493296 | 3485583 | 3488197 | 3481894 | 3484281 | 3459282 | 3480433 | 3483309 | 3483774 | 3483774 | 3482943 | 3480466 | 3466048 | 3462207 | 3486048 | 3486355 | 3488863 | 3483245 | 3483826 | 3480406 | 3481812 | 3483360 | 3479799 | 3466880 | 3476950 | 3484446 | 3479912 | 3477020 | 3479779 | 3481633 | 3482311 | 3479665 | 3478316 | 3476983 | 3477007 | 3478487 | 3478367 | 3476565 | 3475035 | 3476266 | 3457008 | 3471297 | 3475819 | 3468228 | 3475920 | 3480004 | 3475914 | 3474231 | 3477643 | 3472334 | 3470724 | 3476625 | 3477679 | 3472944 | 3478066 | 3478802 | 106 | 3478558.1194 | 9348.8209 | 0.0026876
107 | 3243966 | 3237223 | 3238700 | 3230584 | 3188886 | 3236056 | 3241840 | 3239136 | 3207865 | 3229838 | 3237394 | 3234911 | 3233942 | 3235744 | 3233139 | 3236656 | 3210305 | 3226704 | 3234687 | 3233462 | 3233462 | 3234452 | 3231241 | 3221060 | 3207535 | 3234418 | 3235504 | 3236245 | 3230653 | 3232258 | 3234241 | 3232529 | 3237046 | 3231100 | 3219336 | 3232179 | 3237380 | 3232903 | 3228230 | 3235863 | 3235251 | 3232567 | 3229036 | 3226589 | 3226190 | 3232070 | 3227586 | 3225123 | 3224735 | 3224380 | 3224333 | 3201885 | 3218241 | 3224095 | 3216321 | 3227485 | 3227106 | 3225139 | 3225959 | 3229343 | 3224273 | 3220610 | 3227446 | 3227040 | 3226156 | 3226554 | 3224737 | 107 | 3228491.3881 | 9316.1584 | 0.0028856
108 | 2332628 | 2331728 | 2329916 | 2330626 | 2302264 | 2334472 | 2334402 | 2334293 | 2310250 | 2327513 | 2336771 | 2330584 | 2329193 | 2328735 | 2326750 | 2326699 | 2313165 | 2323505 | 2328425 | 2331110 | 2331110 | 2327575 | 2327840 | 2321611 | 2311814 | 2327856 | 2331188 | 2330204 | 2326694 | 2322362 | 2322700 | 2328620 | 2326178 | 2325374 | 2316907 | 2328573 | 2331013 | 2328504 | 2325034 | 2328989 | 2326334 | 2326809 | 2322593 | 2324454 | 2321398 | 2321184 | 2320408 | 2323203 | 2316327 | 2316564 | 2316470 | 2304116 | 2315760 | 2318527 | 2315452 | 2319704 | 2319967 | 2319199 | 2320707 | 2318299 | 2317383 | 2319869 | 2320209 | 2320721 | 2319251 | 2320605 | 2325623 | 108 | 2323795.6866 | 6926.7909 | 0.0029808
109 | 3227162 | 3225272 | 3226364 | 3217990 | 3175895 | 3222118 | 3225409 | 3223954 | 3197224 | 3216067 | 3222631 | 3225551 | 3221011 | 3218533 | 3222681 | 3221963 | 3195140 | 3214281 | 3221419 | 3225739 | 3225739 | 3222989 | 3223514 | 3209143 | 3195617 | 3222618 | 3221550 | 3224902 | 3221292 | 3220847 | 3220150 | 3221034 | 3222776 | 3222935 | 3205578 | 3218383 | 3220794 | 3225403 | 3213956 | 3218946 | 3218661 | 3222311 | 3216645 | 3212729 | 3215351 | 3221244 | 3216387 | 3213475 | 3212482 | 3211672 | 3213827 | 3192893 | 3212313 | 3212437 | 3207970 | 3213811 | 3215839 | 3215889 | 3216173 | 3216628 | 3215762 | 3213157 | 3218383 | 3213031 | 3212711 | 3217107 | 3216659 | 109 | 3216568.9104 | 8890.8167 | 0.0027641
110 | 3680498 | 3672625 | 3677560 | 3669327 | 3618450 | 3670134 | 3678041 | 3675453 | 3638303 | 3670658 | 3679768 | 3678156 | 3674732 | 3676549 | 3667416 | 3675342 | 3645450 | 3668720 | 3675453 | 3677217 | 3677217 | 3672100 | 3672693 | 3660923 | 3648414 | 3675048 | 3677312 | 3678207 | 3671429 | 3674659 | 3677026 | 3678417 | 3679778 | 3674298 | 3660072 | 3671999 | 3676205 | 3674199 | 3662505 | 3671123 | 3670572 | 3669877 | 3671763 | 3667970 | 3665990 | 3671681 | 3672556 | 3665614 | 3669143 | 3666097 | 3668911 | 3645135 | 3665132 | 3668021 | 3661642 | 3668081 | 3670699 | 3668389 | 3668194 | 3673855 | 3670614 | 3666747 | 3671419 | 3672771 | 3667438 | 3667466 | 3670275 | 110 | 3669246.6866 | 10161.7735 | 0.0027694
111 | 3814560 | 3810496 | 3810852 | 3807611 | 3749471 | 3807955 | 3815153 | 3813711 | 3774042 | 3805523 | 3815311 | 3813512 | 3810033 | 3812935 | 3803684 | 3812212 | 3777646 | 3804389 | 3814006 | 3812361 | 3812361 | 3809432 | 3806876 | 3793167 | 3780152 | 3807261 | 3808063 | 3814031 | 3807357 | 3810829 | 3811780 | 3813443 | 3812737 | 3809397 | 3790888 | 3808840 | 3811464 | 3804919 | 3799603 | 3804869 | 3809632 | 3811256 | 3809205 | 3800956 | 3802175 | 3811518 | 3807390 | 3802036 | 3806192 | 3801929 | 3802536 | 3783211 | 3796406 | 3802225 | 3795057 | 3801732 | 3807935 | 3804531 | 3802550 | 3803959 | 3805190 | 3802669 | 3802828 | 3803776 | 3805026 | 3804356 | 3804179 | 111 | 3804378.9104 | 10854.5991 | 0.0028532
112 | 3493827 | 3488098 | 3490158 | 3481478 | 3432423 | 3488965 | 3489897 | 3491145 | 3454209 | 3485869 | 3490577 | 3489769 | 3481817 | 3486671 | 3486098 | 3484724 | 3461582 | 3478852 | 3487135 | 3488735 | 3488735 | 3488192 | 3487473 | 3473091 | 3461619 | 3485019 | 3484870 | 3488077 | 3487345 | 3490813 | 3487044 | 3490217 | 3489635 | 3488233 | 3471155 | 3485136 | 3484117 | 3483284 | 3481346 | 3484544 | 3482187 | 3482803 | 3484970 | 3485544 | 3479062 | 3483338 | 3478685 | 3478931 | 3475761 | 3476363 | 3476366 | 3458278 | 3472375 | 3483007 | 3471252 | 3478431 | 3484608 | 3479856 | 3481185 | 3478739 | 3481194 | 3479199 | 3478504 | 3479899 | 3480593 | 3479340 | 3481726 | 112 | 3481256.2687 | 9789.9523 | 0.0028122
113 | 3250313 | 3244423 | 3245968 | 3239686 | 3196344 | 3243666 | 3244251 | 3245474 | 3214897 | 3239284 | 3244667 | 3248714 | 3243497 | 3242953 | 3240912 | 3242235 | 3218648 | 3237694 | 3244986 | 3245154 | 3245154 | 3242203 | 3243142 | 3230433 | 3219116 | 3245150 | 3244875 | 3245886 | 3243296 | 3244376 | 3243580 | 3245400 | 3245872 | 3245007 | 3229522 | 3240705 | 3245235 | 3240291 | 3235949 | 3237769 | 3240966 | 3241015 | 3233067 | 3235548 | 3235375 | 3237656 | 3237945 | 3232728 | 3232955 | 3233783 | 3231861 | 3214265 | 3230553 | 3236488 | 3226280 | 3233007 | 3235914 | 3237708 | 3234170 | 3237324 | 3238032 | 3233688 | 3239182 | 3235057 | 3230195 | 3234574 | 3236967 | 113 | 3237507.9104 | 9061.4263 | 0.0027989
114 | 3960074 | 3958472 | 3956548 | 3953781 | 3894576 | 3953076 | 3961995 | 3960564 | 3918252 | 3950917 | 3959008 | 3958977 | 3955114 | 3954830 | 3954683 | 3955629 | 3922514 | 3945809 | 3958083 | 3957518 | 3957518 | 3953295 | 3957313 | 3939598 | 3930069 | 3954211 | 3957248 | 3960170 | 3952610 | 3955387 | 3957745 | 3958697 | 3956429 | 3955907 | 3936040 | 3951542 | 3958462 | 3954804 | 3945839 | 3953475 | 3955843 | 3956698 | 3950935 | 3949706 | 3946872 | 3950848 | 3953222 | 3951263 | 3947690 | 3947739 | 3945565 | 3923756 | 3942746 | 3954206 | 3939819 | 3948488 | 3952642 | 3947038 | 3949622 | 3950658 | 3948324 | 3946972 | 3950849 | 3948116 | 3945621 | 3948020 | 3947780 | 114 | 3949818.1642 | 11036.3401 | 0.0027941
115 | 3194753 | 3192840 | 3189663 | 3187837 | 3149009 | 3195609 | 3197695 | 3196217 | 3164912 | 3189080 | 3193159 | 3195738 | 3190386 | 3194002 | 3191923 | 3189707 | 3167340 | 3186370 | 3191559 | 3194295 | 3194295 | 3191452 | 3189082 | 3178040 | 3168518 | 3190524 | 3191989 | 3191481 | 3189645 | 3189193 | 3191322 | 3192263 | 3193092 | 3191725 | 3175789 | 3187277 | 3190328 | 3189733 | 3182255 | 3188834 | 3185663 | 3187532 | 3182892 | 3183588 | 3184093 | 3186611 | 3181459 | 3181878 | 3178861 | 3181553 | 3176157 | 3161753 | 3176303 | 3181449 | 3174346 | 3179749 | 3183044 | 3184715 | 3182214 | 3183328 | 3184713 | 3180496 | 3182540 | 3181663 | 3180913 | 3182856 | 3185199 | 115 | 3185141.7761 | 8730.5599 | 0.002741
116 | 3379335 | 3377933 | 3375061 | 3369696 | 3323195 | 3371270 | 3376943 | 3379261 | 3344870 | 3369285 | 3379320 | 3378379 | 3374390 | 3375169 | 3373717 | 3372304 | 3346723 | 3366756 | 3376250 | 3375339 | 3375339 | 3373068 | 3373105 | 3358331 | 3348409 | 3373969 | 3373599 | 3375315 | 3371928 | 3376731 | 3375774 | 3381398 | 3377546 | 3377546 | 3361257 | 3372370 | 3376828 | 3374257 | 3367367 | 3372249 | 3371764 | 3373235 | 3369835 | 3368559 | 3364120 | 3369491 | 3364963 | 3365832 | 3363704 | 3366424 | 3358986 | 3344123 | 3361270 | 3365537 | 3355686 | 3362484 | 3368430 | 3370004 | 3364608 | 3370961 | 3369469 | 3362706 | 3365965 | 3366300 | 3368525 | 3366513 | 3370920 | 116 | 3368537.2537 | 9811.5585 | 0.0029127
117 | 2572452 | 2573826 | 2571355 | 2570392 | 2537515 | 2571067 | 2576069 | 2573996 | 2547309 | 2571360 | 2573296 | 2574291 | 2573541 | 2573254 | 2568392 | 2569702 | 2550084 | 2566700 | 2573725 | 2570479 | 2570479 | 2570822 | 2572871 | 2561366 | 2554365 | 2566369 | 2570409 | 2570925 | 2568572 | 2573198 | 2570638 | 2575507 | 2570091 | 2567856 | 2561266 | 2568161 | 2572859 | 2566410 | 2562453 | 2565558 | 2567978 | 2565759 | 2565086 | 2562596 | 2562735 | 2563650 | 2564209 | 2562811 | 2563999 | 2561819 | 2561615 | 2546169 | 2557704 | 2560521 | 2556243 | 2560300 | 2565061 | 2563924 | 2566167 | 2561782 | 2564116 | 2560221 | 2559738 | 2560862 | 2560629 | 2560038 | 2562676 | 117 | 2565483.403 | 7303.7693 | 0.0028469
118 | 4044421 | 4039245 | 4039882 | 4034097 | 3976016 | 4034833 | 4037452 | 4042777 | 4005830 | 4038080 | 4040032 | 4044907 | 4037580 | 4038668 | 4041110 | 4039255 | 4004175 | 4030623 | 4041663 | 4041887 | 4041887 | 4040716 | 4035252 | 4021309 | 4008201 | 4042683 | 4041118 | 4043654 | 4036802 | 4042767 | 4038839 | 4041659 | 4038796 | 4038640 | 4017572 | 4039157 | 4041050 | 4036141 | 4027774 | 4038277 | 4033496 | 4035749 | 4035086 | 4035200 | 4033721 | 4035485 | 4033645 | 4033387 | 4028022 | 4029359 | 4030236 | 4010418 | 4028302 | 4035285 | 4023797 | 4030279 | 4035368 | 4037691 | 4033254 | 4030993 | 4035307 | 4031275 | 4034709 | 4027436 | 4031308 | 4031964 | 4035924 | 118 | 4033157.0597 | 11150.4262 | 0.0027647
119 | 4105278 | 4101909 | 4104836 | 4100555 | 4036228 | 4095128 | 4102836 | 4103917 | 4062247 | 4098826 | 4103660 | 4105089 | 4101846 | 4103227 | 4097822 | 4100029 | 4067992 | 4092178 | 4102313 | 4100133 | 4100133 | 4103641 | 4101617 | 4085259 | 4067735 | 4104145 | 4102219 | 4108589 | 4099246 | 4100708 | 4105069 | 4105939 | 4103646 | 4101821 | 4082642 | 4096158 | 4105885 | 4100173 | 4088697 | 4099113 | 4099905 | 4100791 | 4097737 | 4100382 | 4094093 | 4097969 | 4096052 | 4094128 | 4091557 | 4095827 | 4094778 | 4067769 | 4089906 | 4092614 | 4085966 | 4093951 | 4098547 | 4094268 | 4094403 | 4092735 | 4094817 | 4093433 | 4093523 | 4095302 | 4092790 | 4094338 | 4095941 | 119 | 4095253.8209 | 11690.6054 | 0.0028547
120 | 3971834 | 3971342 | 3966200 | 3962393 | 3907224 | 3964292 | 3971190 | 3968013 | 3930613 | 3965377 | 3973359 | 3965385 | 3964202 | 3968646 | 3962677 | 3966673 | 3935305 | 3959630 | 3969471 | 3965762 | 3965762 | 3968839 | 3961812 | 3950979 | 3936488 | 3966427 | 3966106 | 3972436 | 3964584 | 3965515 | 3969423 | 3970086 | 3968157 | 3962689 | 3950027 | 3960697 | 3967267 | 3962545 | 3960067 | 3963868 | 3966179 | 3962354 | 3963907 | 3962059 | 3961634 | 3964321 | 3959927 | 3957378 | 3960238 | 3959698 | 3954884 | 3935094 | 3953096 | 3956538 | 3954785 | 3957057 | 3962534 | 3963631 | 3960224 | 3960394 | 3959586 | 3957538 | 3960707 | 3960887 | 3956601 | 3961617 | 3963366 | 120 | 3960590.9851 | 10663.0431 | 0.0026923
121 | 3864874 | 3863095 | 3861210 | 3854534 | 3797921 | 3854772 | 3863404 | 3860542 | 3822371 | 3856697 | 3859009 | 3862519 | 3857684 | 3861269 | 3854465 | 3857721 | 3829469 | 3852611 | 3857873 | 3864469 | 3864469 | 3859131 | 3856441 | 3840814 | 3828987 | 3858798 | 3860427 | 3861863 | 3858931 | 3859491 | 3855811 | 3860862 | 3859415 | 3853234 | 3836799 | 3856984 | 3859514 | 3853048 | 3848244 | 3856975 | 3853601 | 3860303 | 3851749 | 3855108 | 3851946 | 3855348 | 3852106 | 3852348 | 3851444 | 3853243 | 3849287 | 3826351 | 3845883 | 3849724 | 3843334 | 3848412 | 3855197 | 3853776 | 3851407 | 3854676 | 3855164 | 3849842 | 3854573 | 3853654 | 3852171 | 3853700 | 3857878 | 121 | 3852820.4776 | 10967.3085 | 0.0028466
122 | 3971998 | 3968874 | 3969513 | 3965811 | 3913286 | 3967585 | 3971717 | 3974338 | 3931404 | 3963764 | 3972439 | 3968664 | 3970591 | 3969108 | 3969163 | 3966707 | 3936415 | 3959422 | 3966932 | 3968152 | 3968152 | 3967426 | 3967610 | 3950356 | 3940607 | 3967569 | 3964311 | 3971576 | 3971129 | 3968222 | 3968260 | 3967427 | 3971717 | 3972806 | 3956966 | 3965208 | 3972325 | 3965527 | 3956962 | 3963351 | 3966610 | 3965533 | 3969185 | 3963687 | 3961390 | 3964235 | 3962955 | 3963333 | 3957308 | 3962096 | 3961358 | 3938893 | 3956264 | 3960386 | 3950931 | 3961834 | 3964318 | 3962862 | 3960397 | 3958597 | 3963595 | 3961092 | 3962325 | 3962859 | 3961412 | 3963638 | 3966238 | 122 | 3962637.6269 | 10305.5007 | 0.0026007
123 | 4044727 | 4046186 | 4047848 | 4038929 | 3981969 | 4047158 | 4049289 | 4041642 | 4006440 | 4037748 | 4050522 | 4048461 | 4044232 | 4046091 | 4043414 | 4042654 | 4008010 | 4041791 | 4049194 | 4047827 | 4047827 | 4045709 | 4044754 | 4029194 | 4015509 | 4045043 | 4047436 | 4048253 | 4048098 | 4042813 | 4041731 | 4047516 | 4046025 | 4042499 | 4025668 | 4040235 | 4045869 | 4043764 | 4035959 | 4043192 | 4043875 | 4045686 | 4041873 | 4046250 | 4038673 | 4041891 | 4039188 | 4034583 | 4037810 | 4035390 | 4038636 | 4013470 | 4031701 | 4038266 | 4028209 | 4036058 | 4040636 | 4038853 | 4041645 | 4038202 | 4039228 | 4037643 | 4038147 | 4041465 | 4039150 | 4038341 | 4043934 | 123 | 4039104.9104 | 11393.2309 | 0.0028207
124 | 4325251 | 4331562 | 4325151 | 4318152 | 4256203 | 4318277 | 4325194 | 4321524 | 4286649 | 4322321 | 4324207 | 4324424 | 4323871 | 4320718 | 4319967 | 4320481 | 4289705 | 4316775 | 4323707 | 4327701 | 4327701 | 4325432 | 4322143 | 4306447 | 4291240 | 4321034 | 4325597 | 4327505 | 4327459 | 4326970 | 4325279 | 4326203 | 4327386 | 4326959 | 4300213 | 4317864 | 4324743 | 4322253 | 4314974 | 4320006 | 4319511 | 4321775 | 4317649 | 4321617 | 4314833 | 4321453 | 4319113 | 4316279 | 4310879 | 4312917 | 4315576 | 4289625 | 4310989 | 4317881 | 4308972 | 4312571 | 4319885 | 4319335 | 4317604 | 4319147 | 4318607 | 4314215 | 4320418 | 4315761 | 4316538 | 4316953 | 4319337 | 124 | 4317293.8507 | 11822.9604 | 0.0027385
125 | 4327750 | 4324671 | 4323485 | 4314202 | 4259837 | 4317492 | 4322174 | 4326034 | 4281392 | 4322610 | 4329799 | 4324779 | 4325909 | 4326882 | 4320807 | 4319324 | 4288002 | 4318172 | 4324110 | 4320169 | 4320169 | 4325201 | 4322020 | 4301948 | 4293676 | 4321794 | 4320955 | 4325199 | 4317615 | 4324855 | 4325427 | 4327133 | 4321688 | 4321545 | 4300655 | 4315899 | 4323607 | 4318590 | 4310379 | 4316956 | 4320051 | 4318403 | 4319095 | 4317649 | 4316386 | 4315944 | 4316115 | 4311114 | 4315130 | 4314636 | 4315602 | 4290797 | 4308646 | 4316739 | 4308717 | 4313463 | 4319577 | 4315813 | 4316250 | 4318207 | 4314635 | 4315191 | 4317256 | 4313125 | 4313849 | 4314152 | 4319987 | 125 | 4315961.791 | 11414.9166 | 0.0026448
126 | 4096551 | 4096249 | 4096454 | 4085563 | 4030460 | 4092339 | 4097012 | 4095293 | 4059864 | 4089239 | 4095073 | 4093957 | 4093582 | 4092008 | 4094627 | 4092180 | 4057537 | 4084684 | 4092240 | 4090807 | 4090807 | 4093247 | 4090484 | 4077720 | 4063049 | 4094305 | 4092081 | 4093407 | 4086828 | 4094627 | 4093333 | 4096664 | 4094611 | 4094412 | 4073334 | 4087015 | 4089631 | 4089996 | 4082833 | 4092032 | 4089395 | 4091738 | 4085796 | 4086586 | 4081733 | 4089508 | 4087290 | 4084039 | 4080976 | 4082977 | 4080909 | 4060070 | 4080216 | 4082622 | 4076965 | 4086457 | 4092015 | 4088110 | 4085543 | 4085181 | 4088790 | 4082608 | 4087271 | 4086157 | 4083455 | 4085902 | 4087426 | 126 | 4086207.0149 | 10969.692 | 0.0026846
127 | 4494949 | 4493603 | 4492717 | 4491012 | 4426631 | 4489189 | 4500170 | 4491284 | 4454576 | 4489324 | 4499758 | 4495929 | 4494077 | 4499030 | 4495845 | 4493991 | 4453357 | 4487125 | 4495171 | 4494848 | 4494848 | 4494116 | 4491646 | 4473385 | 4458483 | 4494677 | 4494921 | 4496624 | 4489654 | 4493852 | 4494566 | 4489983 | 4494280 | 4492222 | 4470582 | 4489375 | 4496091 | 4493244 | 4479942 | 4491591 | 4494012 | 4490906 | 4485568 | 4486891 | 4487478 | 4492828 | 4489684 | 4483001 | 4484089 | 4482627 | 4484927 | 4456862 | 4483473 | 4485138 | 4474119 | 4482061 | 4490427 | 4486499 | 4482990 | 4489210 | 4489518 | 4483666 | 4487600 | 4485562 | 4483317 | 4485534 | 4485190 | 127 | 4486564.8507 | 12351.9189 | 0.0027531
128 | 4517203 | 4510334 | 4512066 | 4503425 | 4445747 | 4504330 | 4515117 | 4510756 | 4466913 | 4505377 | 4509053 | 4515251 | 4510766 | 4512237 | 4506338 | 4507437 | 4471768 | 4501516 | 4513026 | 4513317 | 4513317 | 4507259 | 4505660 | 4491088 | 4477267 | 4506489 | 4507069 | 4513653 | 4507112 | 4510882 | 4509539 | 4510517 | 4508796 | 4505682 | 4487366 | 4503204 | 4511264 | 4508146 | 4497121 | 4506820 | 4504870 | 4507830 | 4506166 | 4507487 | 4502362 | 4506944 | 4505411 | 4504836 | 4496349 | 4500619 | 4498538 | 4476104 | 4495763 | 4500197 | 4486911 | 4496078 | 4502881 | 4499451 | 4501689 | 4502204 | 4501151 | 4500034 | 4505473 | 4498984 | 4500254 | 4502899 | 4502347 | 128 | 4502299.403 | 11990.3777 | 0.0026632
129 | 4545652 | 4539021 | 4542998 | 4530726 | 4468886 | 4531684 | 4539418 | 4538046 | 4493039 | 4530301 | 4538787 | 4539112 | 4533977 | 4540950 | 4533544 | 4534643 | 4499782 | 4529542 | 4536976 | 4537062 | 4537062 | 4536438 | 4535940 | 4513135 | 4506253 | 4532164 | 4539920 | 4535706 | 4533943 | 4536133 | 4539314 | 4540158 | 4538844 | 4536836 | 4516185 | 4529371 | 4537425 | 4536222 | 4523841 | 4533021 | 4534827 | 4533026 | 4530330 | 4532940 | 4530707 | 4533407 | 4529311 | 4524080 | 4526842 | 4525573 | 4526773 | 4498923 | 4522689 | 4528455 | 4519930 | 4526369 | 4530850 | 4529628 | 4527719 | 4530679 | 4531388 | 4530094 | 4528983 | 4530910 | 4533108 | 4531240 | 4531621 | 129 | 4529588.9403 | 12328.7316 | 0.0027218
130 | 3460081 | 3456088 | 3455079 | 3450357 | 3410122 | 3454996 | 3459567 | 3455723 | 3423447 | 3451876 | 3455104 | 3455213 | 3452993 | 3455158 | 3451441 | 3447410 | 3427445 | 3446896 | 3454636 | 3455003 | 3455003 | 3450946 | 3452218 | 3441696 | 3433719 | 3455413 | 3455712 | 3461846 | 3456195 | 3452599 | 3457827 | 3465574 | 3464821 | 3462305 | 3447183 | 3460285 | 3460584 | 3457086 | 3452754 | 3453706 | 3451040 | 3449980 | 3448066 | 3447898 | 3441303 | 3444421 | 3444693 | 3440257 | 3439867 | 3443843 | 3440056 | 3423975 | 3435531 | 3443324 | 3441277 | 3444333 | 3450980 | 3452486 | 3449221 | 3445664 | 3450118 | 3443624 | 3451773 | 3450507 | 3444899 | 3448894 | 3448169 | 130 | 3449138.8955 | 9795.7395 | 0.0028401
131 | 3165328 | 3162457 | 3160902 | 3156192 | 3120346 | 3163839 | 3168953 | 3164968 | 3136140 | 3158372 | 3163962 | 3167457 | 3161030 | 3160426 | 3158124 | 3163128 | 3139807 | 3155024 | 3163800 | 3164397 | 3164397 | 3161979 | 3161072 | 3150417 | 3141336 | 3162525 | 3162520 | 3163951 | 3162282 | 3166292 | 3168009 | 3166861 | 3166739 | 3164493 | 3151581 | 3162952 | 3164528 | 3159598 | 3150884 | 3158565 | 3157676 | 3154100 | 3157590 | 3155107 | 3152108 | 3157035 | 3151596 | 3152004 | 3153027 | 3154078 | 3152124 | 3138444 | 3148919 | 3154032 | 3146261 | 3151921 | 3156443 | 3156602 | 3153900 | 3153879 | 3159615 | 3152296 | 3158803 | 3153444 | 3152980 | 3153424 | 3155664 | 131 | 3156846.3433 | 8386.1796 | 0.0026565
132 | 4530698 | 4527907 | 4526245 | 4516440 | 4460713 | 4524509 | 4529452 | 4524516 | 4485800 | 4516368 | 4529597 | 4527993 | 4526969 | 4526800 | 4521712 | 4525605 | 4486978 | 4518386 | 4525349 | 4525054 | 4525054 | 4525876 | 4520879 | 4500352 | 4493432 | 4523190 | 4526056 | 4529160 | 4519597 | 4521400 | 4523945 | 4527768 | 4524378 | 4527492 | 4509006 | 4520391 | 4524714 | 4520318 | 4515987 | 4519205 | 4516662 | 4518574 | 4518883 | 4514387 | 4518792 | 4522282 | 4516480 | 4514357 | 4509817 | 4512803 | 4509759 | 4488287 | 4511326 | 4514963 | 4506284 | 4514635 | 4520782 | 4518916 | 4516876 | 4513461 | 4517601 | 4516739 | 4516796 | 4513523 | 4513393 | 4515260 | 4513941 | 132 | 4517027.9104 | 11844.8266 | 0.0026223
133 | 4304169 | 4305678 | 4304674 | 4295787 | 4240996 | 4299991 | 4304549 | 4300248 | 4265406 | 4296465 | 4307741 | 4304215 | 4296465 | 4299581 | 4299860 | 4297456 | 4267091 | 4293209 | 4302630 | 4296719 | 4296719 | 4298024 | 4300631 | 4281565 | 4270521 | 4298776 | 4300452 | 4299969 | 4300976 | 4305860 | 4300804 | 4301798 | 4304980 | 4302506 | 4285822 | 4298893 | 4296294 | 4298764 | 4287654 | 4294519 | 4294191 | 4295210 | 4291570 | 4289135 | 4288037 | 4289923 | 4292409 | 4286616 | 4283272 | 4292388 | 4286545 | 4261428 | 4284443 | 4286896 | 4281125 | 4285024 | 4288706 | 4294351 | 4291283 | 4291379 | 4288629 | 4286256 | 4291387 | 4291232 | 4290910 | 4290276 | 4290468 | 133 | 4292411.1343 | 11366.8663 | 0.0026481
134 | 4035397 | 4035347 | 4034382 | 4023225 | 3977325 | 4031689 | 4037421 | 4037856 | 4000992 | 4032802 | 4038768 | 4040772 | 4036534 | 4035701 | 4033071 | 4033216 | 4004834 | 4030357 | 4036610 | 4037108 | 4037108 | 4038753 | 4033447 | 4020523 | 4009017 | 4037195 | 4035234 | 4041716 | 4040910 | 4042863 | 4046413 | 4051172 | 4048492 | 4038667 | 4016350 | 4029845 | 4038528 | 4030293 | 4019999 | 4027097 | 4031847 | 4027226 | 4026522 | 4027580 | 4024511 | 4028782 | 4024151 | 4022216 | 4024171 | 4026621 | 4024674 | 4003330 | 4024630 | 4026960 | 4019170 | 4024360 | 4031278 | 4030857 | 4026071 | 4026272 | 4029221 | 4028760 | 4027258 | 4028946 | 4024338 | 4029574 | 4024788 | 134 | 4029121.5373 | 11416.2842 | 0.0028334
135 | 1087421 | 1089534 | 1089494 | 1087602 | 1073894 | 1088523 | 1088486 | 1091066 | 1079086 | 1087867 | 1090028 | 1091444 | 1090065 | 1090131 | 1090327 | 1088509 | 1078803 | 1086311 | 1091531 | 1090331 | 1090331 | 1090685 | 1088712 | 1084187 | 1082055 | 1089676 | 1089736 | 1090570 | 1090439 | 1089932 | 1090673 | 1091906 | 1089419 | 1088945 | 1085811 | 1086660 | 1090640 | 1086103 | 1087452 | 1088678 | 1089361 | 1090002 | 1087879 | 1088933 | 1090906 | 1087794 | 1087809 | 1087967 | 1088189 | 1086397 | 1088283 | 1082403 | 1087310 | 1088725 | 1084088 | 1086282 | 1088520 | 1089300 | 1090088 | 1090051 | 1087174 | 1087033 | 1085543 | 1088416 | 1088367 | 1087591 | 1087330 | 135 | 1087982.1493 | 3103.2597 | 0.0028523
136 | 1153384 | 1152979 | 1154535 | 1151780 | 1134630 | 1151995 | 1153285 | 1154190 | 1143943 | 1153015 | 1154230 | 1155154 | 1152385 | 1154591 | 1153397 | 1154983 | 1144829 | 1153176 | 1155479 | 1154784 | 1154784 | 1153522 | 1153294 | 1147078 | 1147925 | 1155072 | 1153022 | 1154539 | 1153056 | 1154219 | 1154048 | 1154416 | 1153072 | 1152714 | 1148243 | 1152417 | 1154360 | 1150812 | 1151823 | 1151837 | 1154146 | 1155101 | 1151063 | 1153247 | 1154106 | 1155465 | 1152807 | 1152803 | 1152026 | 1153154 | 1152910 | 1145609 | 1151814 | 1153428 | 1151542 | 1152011 | 1155172 | 1153075 | 1151886 | 1154905 | 1152819 | 1150605 | 1153036 | 1151135 | 1151881 | 1151003 | 1153666 | 136 | 1152349.4328 | 3224.6732 | 0.0027983
137 | 1151849 | 1151433 | 1152915 | 1147818 | 1132283 | 1149421 | 1151853 | 1148879 | 1138374 | 1149686 | 1151476 | 1153887 | 1150037 | 1150313 | 1147378 | 1147977 | 1138760 | 1146300 | 1149384 | 1149357 | 1149357 | 1148880 | 1150583 | 1143047 | 1142910 | 1151188 | 1150645 | 1150824 | 1151579 | 1148530 | 1150907 | 1148043 | 1150344 | 1151051 | 1146299 | 1149996 | 1149553 | 1150361 | 1146021 | 1147340 | 1149519 | 1149827 | 1149114 | 1147669 | 1148422 | 1148409 | 1148445 | 1149284 | 1147378 | 1148336 | 1148968 | 1141003 | 1146667 | 1148677 | 1144410 | 1147197 | 1148433 | 1150320 | 1147632 | 1148986 | 1149069 | 1147751 | 1147736 | 1145853 | 1148838 | 1147387 | 1148731 | 137 | 1148222.3731 | 3425.6288 | 0.0029834
138 | 1240921 | 1238527 | 1241362 | 1238123 | 1220464 | 1238189 | 1241395 | 1240675 | 1228258 | 1239530 | 1240047 | 1241317 | 1242522 | 1240203 | 1240302 | 1240846 | 1230066 | 1237140 | 1240936 | 1239149 | 1239149 | 1239037 | 1241219 | 1235479 | 1231438 | 1239698 | 1241104 | 1242624 | 1242171 | 1240659 | 1239484 | 1241619 | 1241166 | 1239886 | 1236546 | 1241405 | 1239560 | 1239403 | 1239077 | 1239682 | 1241476 | 1241121 | 1242567 | 1239538 | 1241314 | 1241554 | 1242388 | 1241082 | 1239639 | 1240659 | 1240149 | 1232908 | 1239032 | 1240228 | 1236851 | 1238716 | 1242348 | 1240414 | 1241522 | 1240858 | 1239678 | 1241419 | 1239658 | 1237996 | 1239939 | 1240304 | 1241585 | 138 | 1239333.1493 | 3569.4063 | 0.0028801
139 | 1048720 | 1049714 | 1051029 | 1047783 | 1033555 | 1047955 | 1051783 | 1049612 | 1040150 | 1047920 | 1051799 | 1051562 | 1048158 | 1049849 | 1049859 | 1050724 | 1041483 | 1049038 | 1052661 | 1051704 | 1051704 | 1050869 | 1050049 | 1046544 | 1043267 | 1050096 | 1051565 | 1050184 | 1051234 | 1049306 | 1050707 | 1052060 | 1048964 | 1048045 | 1044487 | 1048457 | 1049606 | 1049030 | 1049101 | 1051914 | 1051285 | 1049904 | 1050097 | 1051052 | 1050100 | 1053361 | 1052913 | 1049918 | 1051657 | 1051045 | 1047966 | 1043066 | 1049089 | 1050756 | 1046054 | 1048897 | 1050495 | 1048901 | 1048723 | 1048899 | 1049626 | 1050090 | 1050874 | 1048422 | 1048942 | 1050865 | 1050310 | 139 | 1049187.3731 | 3128.4207 | 0.0029818
140 | 918514 | 918070 | 919749 | 918430 | 906254 | 920036 | 919055 | 920374 | 911054 | 917124 | 918701 | 919305 | 920141 | 919922 | 919643 | 919870 | 912461 | 915363 | 918652 | 918465 | 918465 | 918993 | 917900 | 913623 | 913296 | 920294 | 919349 | 919518 | 916339 | 917066 | 916871 | 917834 | 919124 | 918334 | 913790 | 918290 | 918835 | 917878 | 917409 | 918437 | 915735 | 918063 | 917530 | 919173 | 917934 | 918862 | 918114 | 918135 | 917049 | 917239 | 918812 | 910852 | 915253 | 916889 | 913698 | 916085 | 916308 | 915893 | 917824 | 917711 | 917354 | 915987 | 917542 | 916443 | 915565 | 915140 | 917905 | 140 | 917252.5821 | 2493.2709 | 0.0027182
141 | 1186349 | 1187028 | 1187385 | 1184769 | 1167451 | 1184272 | 1187704 | 1187085 | 1177511 | 1187031 | 1188997 | 1189827 | 1187645 | 1188517 | 1187846 | 1188811 | 1178277 | 1188284 | 1190643 | 1186387 | 1186387 | 1189319 | 1186821 | 1180446 | 1183045 | 1187722 | 1188984 | 1190449 | 1186003 | 1188507 | 1186605 | 1186980 | 1186358 | 1184915 | 1182816 | 1187125 | 1187717 | 1185461 | 1185662 | 1187180 | 1188068 | 1190114 | 1189035 | 1188131 | 1187987 | 1190291 | 1189845 | 1188904 | 1188658 | 1189059 | 1186510 | 1181489 | 1187006 | 1188679 | 1185401 | 1187340 | 1187586 | 1188666 | 1187025 | 1189383 | 1189458 | 1186689 | 1190023 | 1188751 | 1186911 | 1187749 | 1188716 | 141 | 1186832.7612 | 3495.3641 | 0.0029451
142 | 947489 | 946958 | 944375 | 943329 | 932113 | 946907 | 947876 | 945317 | 938026 | 947973 | 945980 | 948213 | 946514 | 946747 | 947210 | 946400 | 937893 | 946354 | 946208 | 947297 | 947297 | 947305 | 945527 | 941265 | 942190 | 947805 | 947853 | 948193 | 947541 | 947012 | 945070 | 946732 | 947894 | 945857 | 942704 | 946598 | 947257 | 946496 | 944796 | 944810 | 947205 | 945933 | 945585 | 946405 | 944894 | 946206 | 946219 | 944758 | 945640 | 945712 | 945299 | 940350 | 944456 | 944651 | 943049 | 943210 | 946259 | 944228 | 946318 | 945443 | 945874 | 945859 | 945835 | 945803 | 945230 | 946745 | 945297 | 142 | 945400.6567 | 2644.6113 | 0.0027973
143 | 783235 | 783532 | 782225 | 780235 | 769276 | 777362 | 778469 | 779048 | 772839 | 780159 | 781547 | 782233 | 780957 | 779859 | 779132 | 780089 | 771064 | 774748 | 778207 | 777597 | 777597 | 778172 | 776764 | 773481 | 771377 | 777240 | 779314 | 778692 | 779040 | 779003 | 778928 | 778543 | 777963 | 779864 | 774573 | 778479 | 781403 | 781493 | 777573 | 779976 | 777410 | 781236 | 778836 | 776852 | 777488 | 776212 | 778360 | 775695 | 777754 | 777512 | 774480 | 769545 | 773199 | 773652 | 771758 | 774213 | 775040 | 774878 | 775398 | 775276 | 774837 | 775042 | 775036 | 774267 | 772738 | 774737 | 775881 | 143 | 777143.5821 | 3189.5347 | 0.0041042
144 | 760099 | 759340 | 758973 | 756735 | 747869 | 759237 | 759934 | 759774 | 754767 | 759060 | 759914 | 759995 | 760912 | 761487 | 760482 | 760641 | 752857 | 759384 | 761574 | 760475 | 760475 | 759254 | 761690 | 757819 | 757096 | 761254 | 761326 | 761336 | 759487 | 760718 | 761465 | 762730 | 761213 | 758567 | 756656 | 760396 | 761224 | 759118 | 761079 | 759278 | 762160 | 762198 | 760050 | 761570 | 761039 | 760934 | 759861 | 759609 | 760800 | 759606 | 760228 | 754037 | 758313 | 760274 | 758674 | 759852 | 761397 | 761993 | 760592 | 761618 | 760938 | 760214 | 761583 | 761050 | 759158 | 760206 | 760176 | 144 | 759758.5075 | 2326.5762 | 0.0030623
145 | 757182 | 756538 | 758505 | 754916 | 744554 | 754338 | 756560 | 755328 | 750201 | 756274 | 755373 | 756949 | 755355 | 756753 | 755209 | 757093 | 750809 | 756722 | 757017 | 759164 | 759164 | 757927 | 755764 | 753078 | 751568 | 757503 | 757958 | 757470 | 758088 | 758264 | 759439 | 758050 | 759462 | 755893 | 754338 | 757273 | 756664 | 757178 | 755428 | 756686 | 757434 | 757238 | 757558 | 757250 | 755924 | 755665 | 756703 | 756544 | 755410 | 755832 | 756741 | 750895 | 756291 | 758404 | 754111 | 756737 | 756685 | 759180 | 755684 | 756782 | 756431 | 756099 | 757263 | 757873 | 756510 | 756625 | 756947 | 145 | 756221.6567 | 2365.8639 | 0.0031285
146 | 766471 | 765883 | 765837 | 764866 | 754395 | 763658 | 764260 | 764220 | 757917 | 761290 | 765183 | 765383 | 765172 | 766508 | 766066 | 765545 | 756489 | 764053 | 766668 | 766430 | 766430 | 766287 | 767406 | 762420 | 759332 | 764736 | 766618 | 766258 | 764665 | 764150 | 765516 | 767364 | 765477 | 764144 | 759286 | 766220 | 766438 | 765079 | 764437 | 765884 | 766152 | 765888 | 766387 | 765907 | 765949 | 765796 | 765912 | 766002 | 765174 | 765319 | 766023 | 759404 | 765013 | 765924 | 763870 | 766189 | 766894 | 766659 | 767046 | 766057 | 765228 | 765078 | 766889 | 764532 | 767742 | 766621 | 767806 | 146 | 764923.9104 | 2562.6195 | 0.0033502
147 | 879099 | 881567 | 880385 | 877777 | 865097 | 878431 | 879464 | 879087 | 872118 | 880216 | 881071 | 880681 | 879906 | 882583 | 880645 | 880776 | 871890 | 879491 | 880709 | 883011 | 883011 | 882102 | 880650 | 875928 | 874668 | 881556 | 881001 | 881999 | 880094 | 880916 | 881461 | 882576 | 881690 | 880572 | 875156 | 881472 | 882883 | 880894 | 880266 | 880060 | 882128 | 882515 | 881683 | 882895 | 881632 | 882193 | 880659 | 879472 | 881035 | 882857 | 882260 | 872633 | 879809 | 882213 | 877494 | 881027 | 881460 | 883967 | 882057 | 880541 | 880909 | 880581 | 881151 | 880769 | 881296 | 882150 | 879151 | 147 | 880141.7313 | 3084.2937 | 0.0035043
148 | 801031 | 803205 | 804705 | 801791 | 788518 | 800513 | 803307 | 803442 | 796003 | 801398 | 803040 | 803523 | 803909 | 805018 | 803610 | 804435 | 794493 | 800708 | 804963 | 803986 | 803986 | 804481 | 801110 | 798609 | 797148 | 803689 | 803710 | 803548 | 803860 | 803229 | 801560 | 803309 | 802980 | 801033 | 797933 | 804083 | 803157 | 803035 | 801910 | 804827 | 803837 | 804942 | 803158 | 803889 | 803312 | 802871 | 802971 | 801818 | 801881 | 803221 | 803610 | 796789 | 803152 | 802480 | 798381 | 802622 | 802706 | 802363 | 800996 | 803814 | 800895 | 802147 | 802250 | 803739 | 801873 | 802710 | 802208 | 148 | 802140.7463 | 2728.0037 | 0.0034009
149 | 925817 | 924806 | 925613 | 924427 | 910828 | 923959 | 926714 | 926726 | 915652 | 925648 | 926189 | 926213 | 928019 | 925248 | 925571 | 925133 | 919451 | 924354 | 926190 | 925566 | 925566 | 927568 | 927530 | 921330 | 922561 | 928093 | 928341 | 927054 | 925810 | 926260 | 927440 | 925864 | 926338 | 924722 | 922278 | 926124 | 925929 | 925584 | 926174 | 928982 | 927424 | 927951 | 928542 | 927061 | 926621 | 928924 | 927947 | 926710 | 925070 | 927706 | 926633 | 919441 | 925740 | 927476 | 923134 | 928151 | 927352 | 927333 | 926791 | 926877 | 928106 | 927185 | 926728 | 925530 | 926667 | 926785 | 928598 | 149 | 925733.6567 | 2937.7965 | 0.0031735
150 | 939276 | 940315 | 939974 | 936955 | 925347 | 938092 | 939945 | 939369 | 930743 | 939468 | 940614 | 940140 | 939179 | 940803 | 939418 | 941846 | 930705 | 939051 | 940444 | 941779 | 941779 | 941833 | 941330 | 936874 | 933738 | 940044 | 939892 | 943106 | 941237 | 939940 | 940318 | 939322 | 939400 | 938598 | 935413 | 940554 | 941540 | 940910 | 938852 | 940564 | 941479 | 942524 | 943255 | 941828 | 939745 | 942385 | 940683 | 941206 | 940640 | 940196 | 939492 | 934709 | 939189 | 940203 | 938134 | 940310 | 941614 | 940141 | 940103 | 940445 | 941465 | 940544 | 940030 | 940747 | 938372 | 940203 | 942604 | 150 | 939566.4627 | 2927.3483 | 0.0031156
151 | 1366429 | 1367943 | 1371107 | 1366316 | 1347221 | 1360421 | 1365857 | 1361753 | 1355505 | 1366729 | 1369764 | 1370283 | 1369933 | 1368906 | 1366077 | 1367197 | 1357613 | 1366858 | 1368558 | 1370466 | 1370466 | 1367739 | 1367201 | 1359612 | 1357101 | 1368802 | 1367650 | 1372378 | 1367528 | 1368150 | 1364458 | 1369374 | 1370055 | 1369203 | 1361585 | 1367188 | 1371062 | 1367068 | 1364224 | 1368909 | 1371881 | 1373166 | 1372437 | 1371311 | 1368914 | 1368192 | 1368421 | 1368634 | 1369363 | 1369297 | 1367044 | 1359708 | 1367219 | 1367227 | 1366459 | 1370378 | 1372668 | 1372452 | 1370347 | 1370628 | 1373711 | 1373825 | 1374031 | 1375544 | 1374225 | 1375367 | 1373184 | 151 | 1367915.2537 | 4934.2921 | 0.0036072
152 | 1396429 | 1396700 | 1400470 | 1395480 | 1378872 | 1396720 | 1399070 | 1398413 | 1386362 | 1397627 | 1400104 | 1400706 | 1396976 | 1397749 | 1398205 | 1399246 | 1388551 | 1395956 | 1398760 | 1398784 | 1398784 | 1400035 | 1398134 | 1394984 | 1391237 | 1401249 | 1400230 | 1400954 | 1398742 | 1400989 | 1400577 | 1401231 | 1400627 | 1398777 | 1391902 | 1396641 | 1400799 | 1398697 | 1396958 | 1398646 | 1398984 | 1401279 | 1400257 | 1399880 | 1395970 | 1399429 | 1398474 | 1397399 | 1396159 | 1398358 | 1397487 | 1388549 | 1394649 | 1399576 | 1395595 | 1396446 | 1398280 | 1398878 | 1398252 | 1395080 | 1398129 | 1396187 | 1398210 | 1397498 | 1396838 | 1398473 | 1399353 | 152 | 1397388.6866 | 3726.5494 | 0.0026668
153 | 1648398 | 1651212 | 1650108 | 1646634 | 1623777 | 1645839 | 1651446 | 1648813 | 1636438 | 1646759 | 1651240 | 1652382 | 1650883 | 1649136 | 1650908 | 1649427 | 1636910 | 1648599 | 1648492 | 1651895 | 1651895 | 1651496 | 1652121 | 1644035 | 1641619 | 1650666 | 1650915 | 1654180 | 1653028 | 1653195 | 1652720 | 1654576 | 1651637 | 1650626 | 1641603 | 1649474 | 1651044 | 1651753 | 1649029 | 1652438 | 1654471 | 1653347 | 1650436 | 1651662 | 1651696 | 1653365 | 1653748 | 1651900 | 1651029 | 1651083 | 1651681 | 1641662 | 1649324 | 1649472 | 1646678 | 1647800 | 1653321 | 1652003 | 1651697 | 1651290 | 1651047 | 1650735 | 1652451 | 1653641 | 1652019 | 1653708 | 1653393 | 153 | 1649731.4179 | 4882.2619 | 0.0029594
154 | 1432505 | 1435247 | 1436260 | 1432736 | 1414621 | 1434081 | 1435686 | 1433694 | 1421088 | 1432366 | 1435333 | 1436731 | 1437057 | 1435930 | 1435807 | 1435931 | 1424102 | 1433911 | 1434493 | 1436227 | 1436227 | 1437926 | 1435286 | 1429711 | 1427652 | 1437756 | 1436732 | 1440119 | 1436447 | 1438035 | 1438502 | 1437866 | 1437788 | 1436914 | 1429934 | 1435740 | 1436192 | 1435662 | 1432773 | 1438416 | 1435257 | 1437621 | 1433336 | 1436817 | 1435197 | 1438840 | 1436129 | 1434144 | 1433635 | 1436367 | 1434438 | 1427462 | 1433472 | 1434656 | 1430840 | 1435731 | 1437896 | 1434802 | 1436392 | 1434347 | 1436922 | 1436343 | 1436499 | 1436487 | 1438415 | 1434804 | 1435892 | 154 | 1434660.0746 | 4082.2447 | 0.0028454
155 | 1282090 | 1281708 | 1281019 | 1275987 | 1260464 | 1278395 | 1281049 | 1283221 | 1269730 | 1279039 | 1281439 | 1283390 | 1282465 | 1282945 | 1284427 | 1283029 | 1269006 | 1279134 | 1283585 | 1283800 | 1283800 | 1283320 | 1281370 | 1275331 | 1273285 | 1282540 | 1282630 | 1283133 | 1284838 | 1282008 | 1279920 | 1281331 | 1278831 | 1279962 | 1273478 | 1280865 | 1284101 | 1283288 | 1278232 | 1281660 | 1284097 | 1285056 | 1282918 | 1284278 | 1283787 | 1283489 | 1280762 | 1281259 | 1282647 | 1283999 | 1281538 | 1273080 | 1280157 | 1280763 | 1277706 | 1281281 | 1283672 | 1283958 | 1282530 | 1280843 | 1280871 | 1280702 | 1281180 | 1281459 | 1280272 | 1283463 | 1282390 | 155 | 1280716.4478 | 4161.5628 | 0.0032494
156 | 1399776 | 1402330 | 1402591 | 1397670 | 1377848 | 1398031 | 1400796 | 1398929 | 1388221 | 1399976 | 1402575 | 1401726 | 1399761 | 1401490 | 1398917 | 1401372 | 1388981 | 1398267 | 1401294 | 1403186 | 1403186 | 1402565 | 1400832 | 1392765 | 1392539 | 1403326 | 1403428 | 1402099 | 1400162 | 1399120 | 1400725 | 1400148 | 1400209 | 1398548 | 1391958 | 1399072 | 1402326 | 1400145 | 1398460 | 1401286 | 1401689 | 1403162 | 1402984 | 1402723 | 1401030 | 1403787 | 1403086 | 1399488 | 1402018 | 1399690 | 1401147 | 1391077 | 1399145 | 1402859 | 1394947 | 1399673 | 1401758 | 1400961 | 1402479 | 1400027 | 1402145 | 1401466 | 1401944 | 1402288 | 1401730 | 1401945 | 1403603 | 156 | 1399783.3881 | 4322.1295 | 0.0030877
157 | 1367870 | 1365549 | 1369070 | 1365208 | 1347122 | 1363825 | 1367957 | 1368160 | 1355989 | 1365100 | 1370193 | 1369763 | 1367141 | 1366825 | 1368794 | 1367901 | 1355338 | 1364753 | 1370562 | 1371097 | 1371097 | 1367901 | 1366983 | 1361060 | 1357808 | 1368026 | 1366449 | 1369131 | 1366415 | 1367420 | 1367173 | 1366723 | 1363479 | 1365824 | 1361266 | 1366687 | 1368449 | 1365558 | 1364425 | 1367991 | 1365804 | 1368646 | 1367076 | 1369171 | 1367654 | 1370007 | 1370787 | 1368637 | 1367675 | 1367614 | 1366552 | 1359295 | 1366529 | 1367317 | 1362086 | 1366644 | 1368090 | 1367038 | 1367328 | 1366553 | 1368178 | 1367437 | 1367859 | 1366630 | 1364500 | 1367058 | 1367397 | 157 | 1366233.4925 | 3918.7851 | 0.0028683
158 | 1352821 | 1353233 | 1352583 | 1349902 | 1330256 | 1350801 | 1352606 | 1352739 | 1340101 | 1351531 | 1355182 | 1353444 | 1353751 | 1353846 | 1353729 | 1351863 | 1340482 | 1349091 | 1355014 | 1353013 | 1353013 | 1351374 | 1353679 | 1348163 | 1344921 | 1352001 | 1353096 | 1355388 | 1352307 | 1354569 | 1353427 | 1352053 | 1352525 | 1352140 | 1344807 | 1353551 | 1354030 | 1353719 | 1351953 | 1352578 | 1353951 | 1353426 | 1353672 | 1353690 | 1351423 | 1353657 | 1356078 | 1351262 | 1352453 | 1353127 | 1352240 | 1343794 | 1350709 | 1353432 | 1350470 | 1352259 | 1352756 | 1351961 | 1352556 | 1351425 | 1352881 | 1354348 | 1351482 | 1350300 | 1349979 | 1353444 | 1352192 | 158 | 1351555.9552 | 3976.7201 | 0.0029423
159 | 1024319 | 1023722 | 1026099 | 1024812 | 1008776 | 1022925 | 1023741 | 1025723 | 1014002 | 1024052 | 1027298 | 1027306 | 1026059 | 1024684 | 1025498 | 1026207 | 1016580 | 1024124 | 1026800 | 1026858 | 1026858 | 1025642 | 1024481 | 1021283 | 1019672 | 1025107 | 1026290 | 1026943 | 1026247 | 1025135 | 1026366 | 1025986 | 1026141 | 1024695 | 1022320 | 1025294 | 1026438 | 1026387 | 1023906 | 1026655 | 1024995 | 1027754 | 1026496 | 1024584 | 1026132 | 1026633 | 1025497 | 1025117 | 1024500 | 1025752 | 1027608 | 1018081 | 1024865 | 1026330 | 1023901 | 1023521 | 1023656 | 1024065 | 1024155 | 1025675 | 1026424 | 1026132 | 1023774 | 1026087 | 1024816 | 1025430 | 1028504 | 159 | 1024655.4478 | 3111.7524 | 0.0030369
160 | 996883 | 995232 | 994880 | 994251 | 979774 | 992919 | 994479 | 995383 | 988499 | 993296 | 995049 | 995063 | 994974 | 993497 | 995392 | 994230 | 987167 | 993428 | 995709 | 994817 | 994817 | 995758 | 995689 | 991603 | 989294 | 995040 | 995994 | 998450 | 994169 | 995953 | 997179 | 996920 | 997082 | 994693 | 990026 | 994889 | 997371 | 994504 | 992077 | 995917 | 996670 | 997036 | 995890 | 996976 | 994606 | 994880 | 995471 | 994050 | 996940 | 994498 | 996019 | 988445 | 993541 | 996442 | 993334 | 994571 | 995985 | 995624 | 995860 | 995611 | 997919 | 995611 | 994626 | 994604 | 995204 | 994468 | 994995 | 160 | 994510.791 | 2812.3003 | 0.0028278
161 | 1087145 | 1088949 | 1088610 | 1083346 | 1069512 | 1084086 | 1087624 | 1086602 | 1076542 | 1085699 | 1087564 | 1086834 | 1083769 | 1087202 | 1086771 | 1085593 | 1076769 | 1083170 | 1085525 | 1085825 | 1085825 | 1086031 | 1085323 | 1080877 | 1078700 | 1084453 | 1084765 | 1083745 | 1084924 | 1083373 | 1084586 | 1084649 | 1083835 | 1084251 | 1077775 | 1083193 | 1083363 | 1082752 | 1081942 | 1082902 | 1085425 | 1086181 | 1083147 | 1084280 | 1084732 | 1083714 | 1084594 | 1083836 | 1081554 | 1083297 | 1084839 | 1076217 | 1082698 | 1082628 | 1080305 | 1081550 | 1080903 | 1083524 | 1083353 | 1082267 | 1082369 | 1081657 | 1084446 | 1081216 | 1080357 | 1083988 | 1084986 | 161 | 1083469.6119 | 3151.4059 | 0.0029086
162 | 1263366 | 1263385 | 1264676 | 1260845 | 1243239 | 1263536 | 1265705 | 1261188 | 1251395 | 1260949 | 1264887 | 1264787 | 1263241 | 1262051 | 1261596 | 1263446 | 1253822 | 1261965 | 1263398 | 1264356 | 1264356 | 1262812 | 1265653 | 1259005 | 1253702 | 1265449 | 1264071 | 1266249 | 1264683 | 1263680 | 1263237 | 1263742 | 1266502 | 1264081 | 1258306 | 1262660 | 1265324 | 1263792 | 1262846 | 1264013 | 1265522 | 1267510 | 1261420 | 1263113 | 1262481 | 1265650 | 1266257 | 1262677 | 1264291 | 1264868 | 1262185 | 1255527 | 1260344 | 1264939 | 1260108 | 1264263 | 1265216 | 1266021 | 1262191 | 1266179 | 1266487 | 1264413 | 1265600 | 1264167 | 1262481 | 1264553 | 1267194 | 162 | 1262860.4925 | 3895.0258 | 0.0030843
163 | 1057745 | 1059561 | 1059031 | 1053304 | 1040105 | 1053746 | 1057353 | 1056477 | 1046688 | 1057068 | 1058841 | 1060107 | 1057362 | 1058681 | 1057793 | 1058404 | 1049648 | 1056468 | 1057680 | 1057742 | 1057742 | 1057550 | 1056571 | 1050700 | 1049931 | 1059198 | 1057950 | 1060547 | 1059438 | 1057939 | 1058584 | 1060431 | 1061029 | 1059504 | 1052561 | 1057144 | 1058444 | 1057878 | 1057443 | 1060354 | 1061788 | 1059597 | 1059906 | 1062164 | 1060581 | 1061601 | 1060869 | 1060870 | 1061459 | 1061024 | 1060411 | 1050611 | 1064612 | 1064289 | 1060711 | 1064308 | 1063772 | 1065070 | 1063891 | 1064602 | 1065638 | 1064592 | 1067307 | 1066730 | 1067533 | 1066537 | 1070960 | 163 | 1059196.6418 | 5047.0566 | 0.004765
164 | 1148673 | 1148805 | 1150234 | 1146581 | 1128750 | 1146400 | 1149652 | 1147727 | 1137455 | 1147170 | 1148138 | 1148794 | 1149090 | 1149245 | 1146571 | 1148559 | 1137142 | 1148151 | 1150127 | 1148346 | 1148346 | 1150255 | 1148629 | 1142156 | 1141579 | 1148812 | 1149172 | 1150572 | 1148962 | 1147321 | 1148306 | 1148979 | 1150547 | 1147585 | 1141578 | 1147706 | 1151027 | 1148234 | 1145494 | 1147734 | 1150002 | 1150327 | 1147746 | 1149350 | 1147995 | 1147909 | 1147682 | 1149928 | 1149588 | 1148109 | 1147718 | 1138083 | 1146165 | 1146911 | 1144529 | 1148763 | 1148635 | 1149429 | 1148337 | 1146152 | 1149564 | 1147341 | 1148672 | 1150173 | 1148503 | 1149280 | 1148600 | 164 | 1147374.5522 | 3671.5866 | 0.0032
165 | 1103993 | 1106459 | 1105746 | 1104165 | 1087740 | 1106333 | 1106903 | 1107176 | 1095327 | 1104983 | 1107935 | 1104905 | 1105846 | 1104368 | 1105543 | 1102636 | 1095319 | 1103748 | 1105468 | 1107383 | 1107383 | 1105485 | 1104259 | 1099968 | 1097614 | 1107369 | 1106852 | 1105549 | 1104550 | 1106686 | 1107452 | 1107630 | 1106458 | 1105004 | 1101473 | 1105241 | 1106828 | 1106141 | 1104532 | 1104859 | 1106384 | 1107098 | 1104845 | 1105684 | 1106837 | 1106780 | 1106645 | 1105940 | 1105959 | 1105611 | 1105396 | 1095690 | 1103859 | 1104767 | 1101576 | 1103544 | 1103675 | 1105033 | 1105072 | 1106649 | 1105225 | 1105282 | 1105642 | 1106236 | 1104905 | 1104871 | 1107674 | 165 | 1104629.6716 | 3409.3371 | 0.0030864
166 | 1182411 | 1183586 | 1185233 | 1180002 | 1162602 | 1181089 | 1182053 | 1182213 | 1170766 | 1181669 | 1183256 | 1181051 | 1182946 | 1182837 | 1180403 | 1183929 | 1171063 | 1179612 | 1185028 | 1184615 | 1184615 | 1183772 | 1183193 | 1176423 | 1175013 | 1184614 | 1184328 | 1185522 | 1181736 | 1181761 | 1183395 | 1185209 | 1185748 | 1181639 | 1178111 | 1182872 | 1185770 | 1180978 | 1181842 | 1184198 | 1183009 | 1184240 | 1182605 | 1181621 | 1182355 | 1183653 | 1185099 | 1182678 | 1182305 | 1182270 | 1181103 | 1173495 | 1179838 | 1183854 | 1180543 | 1182121 | 1182692 | 1183379 | 1183316 | 1182389 | 1184690 | 1182777 | 1183438 | 1184143 | 1183348 | 1184070 | 1182066 | 166 | 1181853.7313 | 3819.1561 | 0.0032315
167 | 1032214 | 1030353 | 1032722 | 1028115 | 1015970 | 1032643 | 1031945 | 1032963 | 1020809 | 1031612 | 1032824 | 1032111 | 1032338 | 1033211 | 1032297 | 1032228 | 1023136 | 1031034 | 1031610 | 1031599 | 1031599 | 1031959 | 1032067 | 1027038 | 1024856 | 1033612 | 1032477 | 1034676 | 1032121 | 1033480 | 1031937 | 1033448 | 1031679 | 1032266 | 1027049 | 1032613 | 1032071 | 1029535 | 1032511 | 1032904 | 1031113 | 1034459 | 1032780 | 1031876 | 1030391 | 1032083 | 1031543 | 1030140 | 1031805 | 1032368 | 1032072 | 1024580 | 1030498 | 1032572 | 1028418 | 1031086 | 1033968 | 1031490 | 1030097 | 1032074 | 1032290 | 1031627 | 1030207 | 1031767 | 1032573 | 1033201 | 1030927 | 167 | 1031039.3582 | 3100.7592 | 0.0030074
168 | 1179122 | 1181185 | 1180719 | 1179185 | 1162289 | 1178064 | 1180268 | 1178772 | 1167792 | 1178202 | 1180246 | 1182885 | 1179628 | 1181148 | 1179851 | 1179385 | 1169983 | 1178358 | 1177887 | 1180236 | 1180236 | 1181023 | 1180711 | 1172535 | 1171356 | 1182058 | 1180811 | 1182966 | 1180594 | 1181764 | 1181681 | 1182585 | 1180098 | 1179206 | 1174499 | 1182385 | 1179224 | 1179130 | 1175910 | 1179735 | 1177708 | 1179831 | 1181247 | 1179269 | 1180838 | 1179149 | 1179571 | 1179890 | 1178840 | 1180218 | 1179604 | 1172432 | 1178379 | 1179219 | 1175301 | 1179127 | 1180180 | 1179616 | 1179248 | 1181145 | 1179699 | 1178066 | 1179210 | 1181151 | 1179648 | 1180975 | 1180101 | 168 | 1178856.3284 | 3500.6949 | 0.0029696
169 | 1303933 | 1306274 | 1302265 | 1302422 | 1282052 | 1299958 | 1302866 | 1302650 | 1291893 | 1303620 | 1305329 | 1305315 | 1302664 | 1302684 | 1302226 | 1302781 | 1292350 | 1301291 | 1304205 | 1304082 | 1304082 | 1306037 | 1303447 | 1297392 | 1296280 | 1305408 | 1304031 | 1305063 | 1304979 | 1306754 | 1307312 | 1307164 | 1307007 | 1303133 | 1298358 | 1304420 | 1307433 | 1304907 | 1302073 | 1306051 | 1307004 | 1305128 | 1304381 | 1307304 | 1304740 | 1305956 | 1305792 | 1303885 | 1304608 | 1305573 | 1305659 | 1297629 | 1302818 | 1305482 | 1299640 | 1304595 | 1305471 | 1307226 | 1304991 | 1305263 | 1304682 | 1304193 | 1307121 | 1305294 | 1305890 | 1305310 | 1306456 | 169 | 1303496.7463 | 4116.8108 | 0.0031583
170 | 1290260 | 1294226 | 1293554 | 1288627 | 1274299 | 1290173 | 1291774 | 1292351 | 1279999 | 1291456 | 1292999 | 1293953 | 1293136 | 1294082 | 1292641 | 1293476 | 1282981 | 1290791 | 1295010 | 1295294 | 1295294 | 1293983 | 1293077 | 1287324 | 1284098 | 1292493 | 1295704 | 1294714 | 1294179 | 1297126 | 1294656 | 1293993 | 1293043 | 1294400 | 1288032 | 1293245 | 1293446 | 1292702 | 1291763 | 1293348 | 1293292 | 1293653 | 1295416 | 1293368 | 1293373 | 1293919 | 1296273 | 1296335 | 1293353 | 1295148 | 1294045 | 1285328 | 1292418 | 1295160 | 1292381 | 1295607 | 1296194 | 1294213 | 1295171 | 1294369 | 1295088 | 1292811 | 1292873 | 1291296 | 1290496 | 1293056 | 1291542 | 170 | 1292445.9701 | 3814.9025 | 0.0029517
171 | 1136666 | 1138063 | 1139920 | 1133309 | 1118260 | 1133327 | 1136746 | 1135685 | 1125499 | 1136883 | 1137959 | 1136873 | 1137250 | 1138253 | 1136031 | 1137473 | 1127685 | 1134738 | 1139489 | 1137771 | 1137771 | 1137571 | 1138367 | 1129226 | 1126727 | 1135987 | 1137999 | 1138427 | 1138375 | 1136713 | 1136340 | 1138291 | 1136195 | 1132517 | 1127966 | 1133658 | 1136816 | 1134225 | 1132308 | 1134612 | 1135350 | 1135400 | 1133701 | 1137597 | 1136360 | 1137729 | 1136453 | 1135572 | 1135766 | 1136317 | 1133924 | 1127773 | 1132978 | 1135550 | 1131995 | 1136596 | 1134699 | 1138275 | 1136265 | 1137283 | 1136201 | 1135794 | 1135398 | 1135115 | 1134222 | 1135865 | 1135354 | 171 | 1135097.0597 | 3655.2188 | 0.0032202
172 | 1108118 | 1110396 | 1110135 | 1106990 | 1093253 | 1107879 | 1110549 | 1108174 | 1099789 | 1109558 | 1110619 | 1110065 | 1109868 | 1110469 | 1110493 | 1109802 | 1098733 | 1107689 | 1109906 | 1110673 | 1110673 | 1111214 | 1110536 | 1104270 | 1105192 | 1111792 | 1109609 | 1111362 | 1110129 | 1110675 | 1108533 | 1110393 | 1108709 | 1107345 | 1102500 | 1108171 | 1110799 | 1108807 | 1110459 | 1111709 | 1111417 | 1111171 | 1110682 | 1111394 | 1109856 | 1111526 | 1112456 | 1111049 | 1108344 | 1108216 | 1108738 | 1101076 | 1107610 | 1107969 | 1106103 | 1109116 | 1110083 | 1109582 | 1107613 | 1107530 | 1108381 | 1108438 | 1110148 | 1107198 | 1109321 | 1110028 | 1109758 | 172 | 1108669.2239 | 3264.8369 | 0.0029448
173 | 1137077 | 1138178 | 1137250 | 1133368 | 1118713 | 1135261 | 1136618 | 1136007 | 1128442 | 1135502 | 1139162 | 1138385 | 1137930 | 1138277 | 1136144 | 1139237 | 1127504 | 1136992 | 1137354 | 1141807 | 1141807 | 1138065 | 1136459 | 1128668 | 1127327 | 1136637 | 1136044 | 1136949 | 1136006 | 1135524 | 1135568 | 1136576 | 1138050 | 1133563 | 1130578 | 1135665 | 1137840 | 1136076 | 1138946 | 1138800 | 1136686 | 1138641 | 1138537 | 1138548 | 1137985 | 1139533 | 1137902 | 1139965 | 1136965 | 1136443 | 1135694 | 1127296 | 1133430 | 1136239 | 1132393 | 1135170 | 1138039 | 1135202 | 1136232 | 1136405 | 1137679 | 1133863 | 1136565 | 1135875 | 1134104 | 1134889 | 1137955 | 173 | 1135859.5672 | 3700.907 | 0.0032582
174 | 1126738 | 1127838 | 1129172 | 1124151 | 1108687 | 1124581 | 1129291 | 1127501 | 1116967 | 1125753 | 1128744 | 1129792 | 1128165 | 1129864 | 1126744 | 1126737 | 1118491 | 1126791 | 1129617 | 1129487 | 1129487 | 1127676 | 1127658 | 1124163 | 1118405 | 1127089 | 1128625 | 1127425 | 1128614 | 1126365 | 1126644 | 1127605 | 1127268 | 1125472 | 1120816 | 1128866 | 1130498 | 1126089 | 1126340 | 1127118 | 1128959 | 1130141 | 1129030 | 1129850 | 1127330 | 1130482 | 1128622 | 1128689 | 1128000 | 1128349 | 1125590 | 1118423 | 1125925 | 1127305 | 1121667 | 1126293 | 1127335 | 1127631 | 1126278 | 1128189 | 1126052 | 1127611 | 1128128 | 1128097 | 1126310 | 1125590 | 1126481 | 174 | 1126562.5522 | 3603.1822 | 0.0031984
175 | 1438850 | 1437103 | 1439427 | 1436099 | 1417491 | 1436485 | 1438630 | 1436670 | 1423012 | 1436643 | 1436391 | 1439851 | 1435236 | 1435464 | 1436022 | 1436967 | 1426699 | 1433928 | 1437432 | 1436639 | 1436639 | 1436946 | 1439079 | 1429890 | 1426983 | 1438420 | 1438103 | 1436196 | 1437918 | 1433434 | 1434146 | 1435263 | 1438036 | 1436379 | 1427735 | 1434822 | 1438391 | 1437924 | 1433302 | 1435696 | 1437081 | 1436664 | 1436184 | 1436182 | 1436724 | 1437526 | 1436573 | 1434524 | 1436147 | 1436952 | 1436272 | 1426555 | 1434395 | 1433644 | 1430726 | 1434961 | 1435888 | 1435956 | 1435406 | 1437389 | 1436204 | 1434338 | 1435627 | 1434192 | 1431252 | 1433451 | 1434123 | 175 | 1435004.1343 | 3871.2144 | 0.0026977
176 | 1635431 | 1635884 | 1636113 | 1633294 | 1608030 | 1632653 | 1635322 | 1632676 | 1619485 | 1633565 | 1635125 | 1631570 | 1633683 | 1631718 | 1631662 | 1630899 | 1618391 | 1627992 | 1633211 | 1632998 | 1632998 | 1632393 | 1629472 | 1625502 | 1623024 | 1632773 | 1636214 | 1633540 | 1634256 | 1634962 | 1636768 | 1634696 | 1636058 | 1631723 | 1625832 | 1631888 | 1635049 | 1630615 | 1629341 | 1633689 | 1633336 | 1636008 | 1633550 | 1632807 | 1630964 | 1631966 | 1631976 | 1632213 | 1632759 | 1634882 | 1632095 | 1620614 | 1632534 | 1634422 | 1629176 | 1631958 | 1635287 | 1635242 | 1634792 | 1634873 | 1631208 | 1631394 | 1635355 | 1631968 | 1633279 | 1634854 | 1636447 | 176 | 1631887.3731 | 4782.733 | 0.0029308
177 | 1672348 | 1673967 | 1674558 | 1670385 | 1650840 | 1671519 | 1672455 | 1675271 | 1658600 | 1669804 | 1673477 | 1674197 | 1671295 | 1669968 | 1670183 | 1673047 | 1657997 | 1670253 | 1673506 | 1672028 | 1672028 | 1670616 | 1672525 | 1664157 | 1660155 | 1672164 | 1672639 | 1673459 | 1670345 | 1675486 | 1672539 | 1675171 | 1670464 | 1671860 | 1662756 | 1669237 | 1673326 | 1671233 | 1667586 | 1671734 | 1674409 | 1672785 | 1676018 | 1670090 | 1672714 | 1672390 | 1675277 | 1671990 | 1669618 | 1671714 | 1671963 | 1658316 | 1671884 | 1671921 | 1668827 | 1672090 | 1670907 | 1674932 | 1673056 | 1673249 | 1673879 | 1671350 | 1671830 | 1669169 | 1671792 | 1670886 | 1673528 | 177 | 1670772.5672 | 4556.2439 | 0.002727
178 | 1540511 | 1544630 | 1540998 | 1537207 | 1519573 | 1539658 | 1542340 | 1540263 | 1530257 | 1539847 | 1542713 | 1544983 | 1541501 | 1543321 | 1542063 | 1542192 | 1528547 | 1539484 | 1542936 | 1543979 | 1543979 | 1543699 | 1542236 | 1537122 | 1534725 | 1544366 | 1543192 | 1543571 | 1542259 | 1543665 | 1546071 | 1544679 | 1544459 | 1542898 | 1535603 | 1540351 | 1542437 | 1540307 | 1539344 | 1542160 | 1540579 | 1541110 | 1540302 | 1541270 | 1540342 | 1540996 | 1540701 | 1540197 | 1539610 | 1539268 | 1537102 | 1530246 | 1536763 | 1541631 | 1535233 | 1537911 | 1541430 | 1542220 | 1542485 | 1540854 | 1540896 | 1541178 | 1543190 | 1542099 | 1538864 | 1541565 | 1543216 | 178 | 1540378.8657 | 4218.0316 | 0.0027383
179 | 1240568 | 1239295 | 1240416 | 1237780 | 1217233 | 1239054 | 1238986 | 1235000 | 1226920 | 1238193 | 1239769 | 1239587 | 1239317 | 1237532 | 1237585 | 1239039 | 1225642 | 1236339 | 1237479 | 1240680 | 1240680 | 1236192 | 1238737 | 1232614 | 1232340 | 1240972 | 1241093 | 1241146 | 1238268 | 1240162 | 1240079 | 1241650 | 1239772 | 1236556 | 1231794 | 1237626 | 1237966 | 1236277 | 1235689 | 1237483 | 1237417 | 1237320 | 1236693 | 1238459 | 1235860 | 1238245 | 1236984 | 1239045 | 1234368 | 1238096 | 1236884 | 1227850 | 1236494 | 1236092 | 1231569 | 1235285 | 1238067 | 1237334 | 1236458 | 1238741 | 1239638 | 1236466 | 1237099 | 1236847 | 1235067 | 1237362 | 1237710 | 179 | 1236880 | 3946.422 | 0.0031906
180 | 1211320 | 1211617 | 1212873 | 1210378 | 1192720 | 1211632 | 1214662 | 1211010 | 1201113 | 1212890 | 1214669 | 1212138 | 1211142 | 1216880 | 1210668 | 1210417 | 1201811 | 1210773 | 1214034 | 1212440 | 1212440 | 1212961 | 1210179 | 1206632 | 1204436 | 1212374 | 1212322 | 1213165 | 1211283 | 1214120 | 1212623 | 1214089 | 1212730 | 1213678 | 1208726 | 1210435 | 1214593 | 1211300 | 1210323 | 1211557 | 1211604 | 1212560 | 1212807 | 1213134 | 1212483 | 1212701 | 1211671 | 1211005 | 1212925 | 1211910 | 1211604 | 1204678 | 1208977 | 1210451 | 1207851 | 1208373 | 1209766 | 1209683 | 1210423 | 1211364 | 1209930 | 1208400 | 1210727 | 1210662 | 1209182 | 1212913 | 1211122 | 180 | 1210806.8507 | 3511.75 | 0.0029003
181 | 1424944 | 1421767 | 1424452 | 1419500 | 1401596 | 1423396 | 1424049 | 1424733 | 1410589 | 1424605 | 1423778 | 1426549 | 1424983 | 1426931 | 1425317 | 1423251 | 1412372 | 1422248 | 1424575 | 1425275 | 1425275 | 1428714 | 1426236 | 1418551 | 1414067 | 1424650 | 1426916 | 1426912 | 1422697 | 1424599 | 1422169 | 1422731 | 1424275 | 1421969 | 1414559 | 1424312 | 1428421 | 1422934 | 1421813 | 1424861 | 1426583 | 1424447 | 1424453 | 1425810 | 1424394 | 1428675 | 1428040 | 1424870 | 1425057 | 1422946 | 1421309 | 1412243 | 1420219 | 1420371 | 1418117 | 1421218 | 1424703 | 1424240 | 1424887 | 1423947 | 1423004 | 1424468 | 1423614 | 1423120 | 1422849 | 1422808 | 1425710 | 181 | 1422890.6418 | 4510.2541 | 0.0031698
182 | 1592077 | 1587855 | 1587648 | 1583118 | 1561923 | 1583948 | 1588464 | 1586649 | 1573416 | 1586672 | 1588070 | 1588997 | 1588056 | 1588097 | 1587104 | 1588174 | 1574563 | 1586596 | 1591089 | 1588761 | 1588761 | 1591335 | 1589204 | 1582947 | 1578801 | 1588798 | 1589943 | 1592238 | 1589899 | 1591400 | 1588946 | 1591096 | 1589838 | 1588903 | 1580573 | 1585430 | 1589808 | 1588806 | 1587043 | 1588416 | 1587520 | 1588973 | 1588077 | 1587146 | 1585624 | 1588753 | 1589242 | 1588432 | 1586802 | 1589163 | 1588679 | 1579414 | 1587558 | 1587247 | 1582859 | 1589684 | 1590997 | 1585999 | 1583983 | 1587637 | 1587907 | 1585712 | 1585381 | 1586489 | 1588252 | 1587668 | 1588306 | 182 | 1586820.3881 | 4664.6258 | 0.0029396
183 | 954878 | 955327 | 956941 | 952396 | 941061 | 953428 | 955941 | 955811 | 946191 | 954593 | 955461 | 958457 | 954666 | 956941 | 954086 | 955423 | 948801 | 955066 | 955492 | 955707 | 955707 | 955631 | 954082 | 951078 | 949439 | 956144 | 957627 | 958764 | 953478 | 954220 | 953141 | 952953 | 953127 | 953559 | 947545 | 954019 | 953663 | 952565 | 953659 | 955558 | 953927 | 954627 | 954716 | 955216 | 952814 | 957861 | 954963 | 956257 | 954362 | 954883 | 954576 | 946562 | 951190 | 952662 | 950307 | 951928 | 955837 | 953894 | 953378 | 952397 | 954255 | 954031 | 952965 | 953823 | 953471 | 952401 | 952682 | 183 | 953710.6119 | 2882.1539 | 0.003022
184 | 804786 | 804203 | 805162 | 801774 | 791299 | 803775 | 801982 | 804267 | 797010 | 803070 | 805689 | 804789 | 806076 | 805189 | 804957 | 805726 | 797787 | 804141 | 804206 | 805273 | 805273 | 806931 | 807997 | 802215 | 798796 | 804282 | 806819 | 807510 | 804977 | 805047 | 805083 | 804600 | 804622 | 803871 | 801459 | 804148 | 805569 | 804532 | 803731 | 805294 | 806091 | 804769 | 802380 | 806228 | 804589 | 805591 | 804318 | 804191 | 804788 | 804460 | 804611 | 798775 | 804414 | 803785 | 802225 | 804914 | 802631 | 804696 | 804553 | 804316 | 804868 | 804565 | 803641 | 805550 | 804811 | 804457 | 804448 | 184 | 804008.6866 | 2512.7856 | 0.0031253
185 | 859965 | 860197 | 861049 | 857886 | 845008 | 858128 | 859963 | 857373 | 852646 | 859794 | 860652 | 861792 | 860558 | 861347 | 860442 | 860790 | 852575 | 861411 | 861645 | 861672 | 861672 | 859818 | 861230 | 855768 | 855219 | 859507 | 860985 | 860666 | 859307 | 860882 | 859649 | 859180 | 858057 | 860523 | 854403 | 859943 | 863043 | 859406 | 857621 | 861177 | 861814 | 863635 | 862611 | 861211 | 863436 | 861190 | 862173 | 860839 | 863709 | 865106 | 860462 | 855342 | 859397 | 860888 | 857630 | 858466 | 860504 | 862252 | 860965 | 860809 | 859746 | 860987 | 859742 | 860026 | 861257 | 862481 | 861943 | 185 | 859874.1791 | 2986.8477 | 0.0034736
186 | 982977 | 983572 | 983835 | 981919 | 968796 | 982272 | 982790 | 983316 | 974244 | 984747 | 986472 | 985926 | 984584 | 986203 | 984407 | 985307 | 973923 | 982712 | 985005 | 984309 | 984309 | 985081 | 984390 | 978451 | 977973 | 985986 | 983054 | 983026 | 983437 | 983301 | 983395 | 983210 | 983104 | 981220 | 977366 | 982217 | 983923 | 982923 | 981467 | 983644 | 983921 | 983807 | 983980 | 986449 | 983168 | 983901 | 983428 | 981977 | 983718 | 983276 | 980872 | 976674 | 979819 | 984124 | 979474 | 982444 | 982853 | 983729 | 981578 | 983150 | 982294 | 982618 | 982037 | 982527 | 982997 | 984471 | 981768 | 186 | 982535.0299 | 2980.0611 | 0.003033
187 | 1033460 | 1035476 | 1033293 | 1032876 | 1016737 | 1032974 | 1034778 | 1032399 | 1022074 | 1033113 | 1034844 | 1034890 | 1036567 | 1036058 | 1035377 | 1034658 | 1025362 | 1031818 | 1035922 | 1037226 | 1037226 | 1036137 | 1035267 | 1029555 | 1027522 | 1034683 | 1036323 | 1036740 | 1033824 | 1035464 | 1034775 | 1035133 | 1032212 | 1030594 | 1027238 | 1032417 | 1034553 | 1031972 | 1032453 | 1035965 | 1036087 | 1035171 | 1035434 | 1035658 | 1034572 | 1035838 | 1035890 | 1036595 | 1035850 | 1035828 | 1033835 | 1024666 | 1032548 | 1034592 | 1031748 | 1033052 | 1036657 | 1035987 | 1035158 | 1034356 | 1035217 | 1032839 | 1034532 | 1033643 | 1034752 | 1035031 | 1033735 | 187 | 1033570.5373 | 3580.6143 | 0.0034643
188 | 911273 | 911799 | 912793 | 909765 | 895778 | 911527 | 913172 | 911564 | 904869 | 911992 | 910353 | 911780 | 914026 | 914699 | 912724 | 913568 | 902126 | 911341 | 914439 | 912076 | 912076 | 911299 | 910190 | 907858 | 904880 | 912768 | 912091 | 912965 | 913605 | 912659 | 913139 | 911529 | 911686 | 910765 | 908408 | 911741 | 914009 | 913248 | 909205 | 911488 | 913191 | 911720 | 913214 | 912573 | 911652 | 913371 | 912906 | 912069 | 910623 | 910764 | 911085 | 905171 | 908757 | 910724 | 907776 | 910434 | 911300 | 911972 | 912346 | 912670 | 912655 | 909490 | 911278 | 912631 | 911359 | 912371 | 911337 | 188 | 911085.2537 | 2972.6136 | 0.0032627
189 | 1008655 | 1011818 | 1012246 | 1007479 | 994393 | 1008588 | 1009980 | 1010390 | 1001745 | 1010002 | 1012136 | 1012397 | 1012692 | 1011901 | 1011784 | 1010242 | 1002214 | 1008612 | 1013010 | 1012040 | 1012040 | 1011018 | 1009991 | 1006350 | 1005039 | 1011096 | 1012774 | 1013564 | 1011914 | 1010334 | 1010211 | 1009731 | 1010013 | 1009326 | 1001567 | 1006782 | 1009905 | 1010230 | 1007108 | 1008944 | 1011979 | 1008893 | 1009268 | 1010910 | 1009798 | 1012264 | 1012908 | 1010187 | 1011603 | 1010617 | 1013307 | 1004454 | 1008462 | 1010277 | 1005358 | 1007543 | 1010059 | 1009504 | 1009752 | 1009984 | 1011037 | 1007439 | 1009664 | 1008307 | 1007763 | 1009884 | 1008273 | 189 | 1009399.3284 | 3186.348 | 0.0031567
190 | 870471 | 871494 | 870235 | 867940 | 854564 | 868125 | 868997 | 869851 | 862279 | 872175 | 869715 | 871180 | 870418 | 871012 | 869574 | 870921 | 862229 | 869508 | 871856 | 870612 | 870612 | 870459 | 869598 | 864405 | 864692 | 869682 | 870430 | 871674 | 870994 | 870832 | 870886 | 870890 | 870087 | 869069 | 864461 | 868859 | 868878 | 869478 | 869938 | 870288 | 870043 | 869759 | 870113 | 868455 | 868962 | 870907 | 870892 | 869142 | 871131 | 870123 | 869348 | 864280 | 866484 | 870525 | 866091 | 868309 | 870164 | 869402 | 867538 | 870519 | 868219 | 869098 | 868346 | 868564 | 868359 | 868608 | 870654 | 190 | 869006.0149 | 2755.1227 | 0.0031704
191 | 1086212 | 1084335 | 1087247 | 1084030 | 1071040 | 1086445 | 1088456 | 1085999 | 1077804 | 1086746 | 1087471 | 1086627 | 1088062 | 1085912 | 1086621 | 1086027 | 1078241 | 1084187 | 1087961 | 1086473 | 1086473 | 1086917 | 1089375 | 1082873 | 1080713 | 1087052 | 1088680 | 1088219 | 1085763 | 1084106 | 1083850 | 1082712 | 1083024 | 1081996 | 1078023 | 1083867 | 1085733 | 1086703 | 1084120 | 1084813 | 1085448 | 1086108 | 1085768 | 1087057 | 1084826 | 1087332 | 1086441 | 1087399 | 1085379 | 1082879 | 1084117 | 1077201 | 1084285 | 1085637 | 1082581 | 1083307 | 1086177 | 1086125 | 1085759 | 1084982 | 1084616 | 1083634 | 1087349 | 1082132 | 1085402 | 1086130 | 1084499 | 191 | 1084828.0299 | 3036.5215 | 0.0027991
192 | 1248389 | 1248818 | 1248711 | 1244321 | 1233591 | 1248252 | 1249852 | 1249506 | 1238276 | 1248411 | 1247855 | 1250967 | 1250390 | 1248161 | 1247930 | 1250198 | 1238361 | 1246841 | 1247004 | 1248048 | 1248048 | 1247482 | 1248313 | 1243226 | 1240111 | 1247140 | 1246674 | 1248883 | 1248282 | 1247350 | 1246532 | 1248173 | 1247072 | 1243615 | 1242921 | 1247488 | 1247959 | 1246216 | 1246151 | 1249959 | 1249267 | 1249892 | 1252405 | 1249332 | 1249186 | 1251230 | 1248062 | 1246629 | 1247323 | 1244571 | 1246937 | 1237586 | 1244387 | 1245465 | 1241384 | 1244855 | 1246266 | 1245146 | 1244775 | 1246065 | 1246881 | 1247492 | 1246233 | 1244862 | 1245392 | 1244374 | 1247210 | 192 | 1246547.5224 | 3322.4459 | 0.0026653
193 | 1283992 | 1282254 | 1284725 | 1281787 | 1263217 | 1280009 | 1281605 | 1282213 | 1272061 | 1281479 | 1282384 | 1284099 | 1284102 | 1284243 | 1283412 | 1283818 | 1269928 | 1281008 | 1282373 | 1284357 | 1284357 | 1284131 | 1282679 | 1277074 | 1272792 | 1278612 | 1281452 | 1282522 | 1282044 | 1281228 | 1278691 | 1281878 | 1281039 | 1277618 | 1273143 | 1277166 | 1281185 | 1280392 | 1278427 | 1281334 | 1280314 | 1282757 | 1278797 | 1279920 | 1278665 | 1281538 | 1281678 | 1280605 | 1280412 | 1278735 | 1279832 | 1271099 | 1276716 | 1279053 | 1276590 | 1278588 | 1278970 | 1277908 | 1277713 | 1278642 | 1278297 | 1278485 | 1281903 | 1278359 | 1276941 | 1278412 | 1277235 | 193 | 1279686.4776 | 3793.0693 | 0.0029641
194 | 1250334 | 1250320 | 1248842 | 1247227 | 1231140 | 1245603 | 1250191 | 1247235 | 1238981 | 1247707 | 1250451 | 1250941 | 1251786 | 1248726 | 1248763 | 1248649 | 1239632 | 1247838 | 1249134 | 1250422 | 1250422 | 1248783 | 1249142 | 1243114 | 1240973 | 1250426 | 1250183 | 1249458 | 1249230 | 1247538 | 1248155 | 1250698 | 1250859 | 1247064 | 1240600 | 1245973 | 1247716 | 1246830 | 1245743 | 1249939 | 1249279 | 1249742 | 1246183 | 1250386 | 1246126 | 1249413 | 1249453 | 1246081 | 1248404 | 1247214 | 1246383 | 1239426 | 1246359 | 1247813 | 1243515 | 1245496 | 1249498 | 1246899 | 1248449 | 1246913 | 1246641 | 1246967 | 1246030 | 1248616 | 1246996 | 1246251 | 1246533 | 194 | 1247281.1045 | 3438.8858 | 0.0027571
195 | 1083581 | 1083629 | 1084479 | 1080287 | 1065350 | 1080808 | 1083327 | 1081350 | 1071769 | 1082473 | 1086913 | 1085042 | 1083083 | 1084212 | 1082411 | 1084359 | 1073740 | 1081415 | 1085001 | 1083958 | 1083958 | 1085265 | 1083336 | 1079013 | 1077008 | 1083471 | 1083046 | 1085524 | 1084797 | 1084791 | 1084056 | 1083641 | 1082495 | 1081612 | 1073911 | 1080984 | 1081708 | 1078302 | 1079966 | 1084264 | 1084491 | 1083262 | 1081395 | 1082984 | 1082555 | 1085453 | 1082788 | 1085458 | 1082219 | 1081588 | 1082844 | 1074264 | 1081235 | 1083377 | 1078587 | 1081307 | 1083842 | 1082252 | 1082567 | 1082094 | 1081665 | 1082762 | 1082472 | 1081884 | 1081291 | 1084188 | 1082651 | 195 | 1081937.4627 | 3507.4901 | 0.0032419
196 | 1319831 | 1321860 | 1321680 | 1318348 | 1302558 | 1320898 | 1324997 | 1323364 | 1310297 | 1321475 | 1323218 | 1323978 | 1321580 | 1321887 | 1321533 | 1322420 | 1310709 | 1318804 | 1323773 | 1323277 | 1323277 | 1325361 | 1320502 | 1317099 | 1312278 | 1324132 | 1325144 | 1325005 | 1321121 | 1322366 | 1322507 | 1320084 | 1321190 | 1318277 | 1313088 | 1320153 | 1325082 | 1322272 | 1318004 | 1319446 | 1322669 | 1323936 | 1323715 | 1321417 | 1322736 | 1322355 | 1323545 | 1323711 | 1321353 | 1321965 | 1321427 | 1310826 | 1318870 | 1320777 | 1316320 | 1320973 | 1320883 | 1319487 | 1319277 | 1319740 | 1318037 | 1318931 | 1318811 | 1319493 | 1320796 | 1320091 | 1320597 | 196 | 1320382.2836 | 3982.0186 | 0.0030158
197 | 1192557 | 1194566 | 1193804 | 1191980 | 1171819 | 1191876 | 1193324 | 1192009 | 1183537 | 1192601 | 1194313 | 1196002 | 1192969 | 1196620 | 1193659 | 1194301 | 1182270 | 1191444 | 1193736 | 1193157 | 1193157 | 1191897 | 1190900 | 1188137 | 1184950 | 1193645 | 1194344 | 1193528 | 1190531 | 1192153 | 1192195 | 1195009 | 1192200 | 1191585 | 1183155 | 1191724 | 1190822 | 1193709 | 1192260 | 1195473 | 1195593 | 1194571 | 1193366 | 1194705 | 1195641 | 1194643 | 1196075 | 1193817 | 1195256 | 1191887 | 1191942 | 1185455 | 1188092 | 1191203 | 1188654 | 1190752 | 1192325 | 1191572 | 1191701 | 1191685 | 1191263 | 1190038 | 1189932 | 1191508 | 1192330 | 1191851 | 1191978 | 197 | 1191727.6567 | 3827.7787 | 0.003212
198 | 1262837 | 1261576 | 1262074 | 1259518 | 1243054 | 1259085 | 1262227 | 1261929 | 1247117 | 1259205 | 1259469 | 1260153 | 1259909 | 1259521 | 1260294 | 1258866 | 1248698 | 1257731 | 1261127 | 1263203 | 1263203 | 1262989 | 1262276 | 1255171 | 1252343 | 1262526 | 1261489 | 1263651 | 1258556 | 1260697 | 1258929 | 1262373 | 1257365 | 1257301 | 1252061 | 1259441 | 1262952 | 1259343 | 1257164 | 1258514 | 1260056 | 1259558 | 1258468 | 1257834 | 1258110 | 1259426 | 1257942 | 1257134 | 1258836 | 1256085 | 1256613 | 1248831 | 1256599 | 1259175 | 1254780 | 1258151 | 1257480 | 1258007 | 1256611 | 1255513 | 1257392 | 1256544 | 1256496 | 1257863 | 1257302 | 1258265 | 1260492 | 198 | 1258291.0448 | 3806.9803 | 0.0030255